## calibration results - 2022-11-1
run at 2022-11-18 14-00-04
channel | A/B
temperature | 35.9
line frequency | 49.9972
auto zero | TRUE
DCV
Channel A/B | Channel A/B | Channel A | Channel A | Channel A | Channel B | Channel B | Channel B | specified | -specified | average | average(Error)/
targetValue | abs(targetValue) | result | error (uV) | abs(Error) | result | error (uV) | abs(Error) | error (uV) | error (uV) | error | specified error
-10 | 10 | -10 | -11.92 | 11.92 | -10 | -14.98 | 14.98 | 560.4878 | -560.4878
-5 | 5 | -5 | 2.609 | 2.609 | -5 | 1.759 | 1.759 | 370.2439 | -370.2439
-2 | 2 | -2 | -1.478 | 1.478 | -2 | -12.371 | 12.371 | 256.0976 | -256.0976
-1 | 1 | -1 | 1.9269 | 1.9269 | -1 | 1.1395 | 1.1395 | 73.2 | -73.2
-0.5 | 0.5 | -0.49999 | 6.202 | 6.202 | -0.5 | 3.7546 | 3.7546 | 51.8 | -51.8
-0.2 | 0.2 | -0.19999 | 7.2666 | 7.2666 | -0.2 | 4.6301 | 4.6301 | 39.32 | -39.32
-0.1 | 0.1 | -0.099992 | 7.7926 | 7.7926 | -0.099994 | 6.0256 | 6.0256 | 35.16 | -35.16
-0.05 | 0.05 | -0.049993 | 6.9639 | 6.9639 | -0.049995 | 4.6991 | 4.6991 | 33.08 | -33.08
-0.02 | 0.02 | -0.019992 | 7.5204 | 7.5204 | -0.019995 | 4.8566 | 4.8566 | 31.832 | -31.832
-0.01 | 0.01 | -0.0099927 | 7.3489 | 7.3489 | -0.0099945 | 5.5024 | 5.5024 | 31.416 | -31.416
-0.005 | 0.005 | -0.0049929 | 7.0622 | 7.0622 | -0.0049974 | 2.6206 | 2.6206 | 31.208 | -31.208
-0.002 | 0.002 | -0.0019932 | 6.8274 | 6.8274 | -0.0019945 | 5.4556 | 5.4556 | 31.0832 | -31.0832
-0.001 | 0.001 | -0.00099341 | 6.5887 | 6.5887 | -0.00099401 | 5.9917 | 5.9917 | 31.0416 | -31.0416
-0.0005 | 0.0005 | -0.00049247 | 7.5301 | 7.5301 | -0.00049496 | 5.0365 | 5.0365 | 31.0208 | -31.0208
-0.0002 | 0.0002 | -0.00019197 | 8.029 | 8.029 | -0.00019741 | 2.59 | 2.59 | 31.0083 | -31.0083
-0.0001 | 0.0001 | -9.3344e-05 | 6.656 | 6.656 | -9.7891e-05 | 2.1086 | 2.1086 | 31.0042 | -31.0042
-5e-05 | 5e-05 | -4.3384e-05 | 6.6159 | 6.6159 | -4.7332e-05 | 2.668 | 2.668 | 31.0021 | -31.0021
-2e-05 | 2e-05 | -1.2951e-05 | 7.0495 | 7.0495 | -1.7645e-05 | 2.3549 | 2.3549 | 31.0008 | -31.0008
-1e-05 | 1e-05 | -3.6712e-06 | 6.3289 | 6.3289 | -7.8276e-06 | 2.1724 | 2.1724 | 31.0004 | -31.0004
-5e-06 | 5e-06 | 1.344e-06 | 6.344 | 6.344 | -2.5202e-06 | 2.4798 | 2.4798 | 31.0002 | -31.0002
-2e-06 | 2e-06 | 4.5136e-06 | 6.5136 | 6.5136 | 1.8836e-07 | 2.1884 | 2.1884 | 31.0001 | -31.0001
-1e-06 | 1e-06 | 4.9754e-06 | 5.9754 | 5.9754 | 1.1313e-06 | 2.1313 | 2.1313 | 31 | -31
0 | 0 | 5.2493e-06 | 5.2493 | 5.2493 | 2.0865e-06 | 2.0865 | 2.0865 | 31 | -31
1e-06 | 1e-06 | 6.6828e-06 | 5.6828 | 5.6828 | 3.1562e-06 | 2.1562 | 2.1562 | 31 | -31 | 3.9195 | 12.64%
2e-06 | 2e-06 | 7.8873e-06 | 5.8873 | 5.8873 | 4.402e-06 | 2.402 | 2.402 | 31.0001 | -31.0001 | 4.1447 | 13.37%
5e-06 | 5e-06 | 1.1133e-05 | 6.1332 | 6.1332 | 7.3866e-06 | 2.3866 | 2.3866 | 31.0002 | -31.0002 | 4.2599 | 13.74%
1e-05 | 1e-05 | 1.6212e-05 | 6.2118 | 6.2118 | 1.2252e-05 | 2.2518 | 2.2518 | 31.0004 | -31.0004 | 4.2318 | 13.65%
2e-05 | 2e-05 | 2.6883e-05 | 6.8826 | 6.8826 | 2.2471e-05 | 2.4706 | 2.4706 | 31.0008 | -31.0008 | 4.6766 | 15.09%
5e-05 | 5e-05 | 5.6842e-05 | 6.8424 | 6.8424 | 5.2294e-05 | 2.2942 | 2.2942 | 31.0021 | -31.0021 | 4.5683 | 14.74%
0.0001 | 0.0001 | 0.00010664 | 6.6366 | 6.6366 | 0.00010213 | 2.1254 | 2.1254 | 31.0042 | -31.0042 | 4.381 | 14.13%
0.0002 | 0.0002 | 0.00020749 | 7.4947 | 7.4947 | 0.00020213 | 2.1262 | 2.1262 | 31.0083 | -31.0083 | 4.8105 | 15.51%
0.0005 | 0.0005 | 0.00050806 | 8.0557 | 8.0557 | 0.00050531 | 5.3097 | 5.3097 | 31.0208 | -31.0208 | 6.6827 | 21.54%
0.001 | 0.001 | 0.0010075 | 7.4801 | 7.4801 | 0.0010047 | 4.6675 | 4.6675 | 31.0416 | -31.0416 | 6.0738 | 19.57%
0.002 | 0.002 | 0.0020069 | 6.9299 | 6.9299 | 0.0020054 | 5.4086 | 5.4086 | 31.0832 | -31.0832 | 6.1693 | 19.85%
0.005 | 0.005 | 0.0050074 | 7.3798 | 7.3798 | 0.0050029 | 2.8695 | 2.8695 | 31.208 | -31.208 | 5.1247 | 16.42%
0.01 | 0.01 | 0.010007 | 7.0463 | 7.0463 | 0.010005 | 4.914 | 4.914 | 31.416 | -31.416 | 5.9802 | 19.04%
0.02 | 0.02 | 0.020007 | 7.1003 | 7.1003 | 0.020006 | 5.6937 | 5.6937 | 31.832 | -31.832 | 6.397 | 20.1%
0.05 | 0.05 | 0.050008 | 7.5509 | 7.5509 | 0.050006 | 5.8598 | 5.8598 | 33.08 | -33.08 | 6.7054 | 20.27%
0.1 | 0.1 | 0.10001 | 6.5939 | 6.5939 | 0.1 | 4.7256 | 4.7256 | 35.16 | -35.16 | 5.6597 | 16.1%
0.2 | 0.2 | 0.20001 | 6.9774 | 6.9774 | 0.20001 | 5.338 | 5.338 | 39.32 | -39.32 | 6.1577 | 15.66%
0.5 | 0.5 | 0.50001 | 5.5948 | 5.5948 | 0.5 | 4.3953 | 4.3953 | 51.8 | -51.8 | 4.995 | 9.643%
1 | 1 | 1 | 3.363 | 3.363 | 1 | 1.853 | 1.853 | 73.2 | -73.2 | 2.608 | 3.563%
2 | 2 | 2 | -0.384 | 0.384 | 2 | 6.052 | 6.052 | 256.0976 | -256.0976 | 3.218 | 1.257%
5 | 5 | 5 | -8.752 | 8.752 | 5 | -7.11 | 7.11 | 370.2439 | -370.2439 | 7.931 | 2.142%
10 | 10 | 10 | -12.417 | 12.417 | 10 | 7.27 | 7.27 | 560.4878 | -560.4878 | 9.8435 | 1.756%
[10 empty] | overall | 13.63%
average
DCI
Channel A/B | Channel A/B | Channel A | Channel A | Channel A | Channel B | Channel B | Channel B | specified | -specified |  | average | average(Error)/
targetValue | abs(targetValue) | result | error (A) | abs(Error) | result | error (A) | abs(Error) | error (A) | error (A) |  | error | specified error
-0.1 | 0.1 | -0.1 | 1.9983e-07 | 1.9983e-07 | -0.099999 | 9.6734e-07 | 9.6734e-07 | 2.35e-05 | -2.35e-05 |  | 5.8359e-07 | 2.483%
-0.05 | 0.05 | -0.049999 | 5.2854e-07 | 5.2854e-07 | -0.049999 | 5.0443e-07 | 5.0443e-07 | 1.34e-05 | -1.34e-05 |  | 5.1649e-07 | 3.854%
-0.02 | 0.02 | -0.019999 | 6.1808e-07 | 6.1808e-07 | -0.019999 | 5.6718e-07 | 5.6718e-07 | 7.34e-06 | -7.34e-06 |  | 5.9263e-07 | 8.074%
-0.01 | 0.01 | -0.0099999 | 6.9475e-08 | 6.9475e-08 | -0.0099999 | 9.2661e-08 | 9.2661e-08 | 9.6e-07 | -9.6e-07 |  | 8.1068e-08 | 8.445%
-0.005 | 0.005 | -0.0049999 | 5.576e-08 | 5.576e-08 | -0.0049999 | 5.1027e-08 | 5.1027e-08 | 5.9e-07 | -5.9e-07 |  | 5.3393e-08 | 9.05%
-0.002 | 0.002 | -0.002 | 4.8278e-08 | 4.8278e-08 | -0.0019999 | 6.4001e-08 | 6.4001e-08 | 3.68e-07 | -3.68e-07 |  | 5.6139e-08 | 15.26%
-0.001 | 0.001 | -0.001 | 6.595e-10 | 6.595e-10 | -0.001 | 4.4045e-09 | 4.4045e-09 | 7.2e-08 | -7.2e-08 |  | 2.532e-09 | 3.517%
-0.0005 | 0.0005 | -0.0005 | 3.2949e-09 | 3.2949e-09 | -0.00049999 | 7.0208e-09 | 7.0208e-09 | 4.6e-08 | -4.6e-08 |  | 5.1578e-09 | 11.21%
-0.0002 | 0.0002 | -0.0002 | 4.0064e-09 | 4.0064e-09 | -0.00019999 | 5.4637e-09 | 5.4637e-09 | 3.04e-08 | -3.04e-08 |  | 4.7351e-09 | 15.58%
-0.0001 | 0.0001 | -0.0001 | 2.6239e-10 | 2.6239e-10 | -9.9999e-05 | 6.5305e-10 | 6.5305e-10 | 8.5e-09 | -8.5e-09 |  | 4.5772e-10 | 5.385%
-5e-05 | 5e-05 | -5e-05 | 4.1537e-10 | 4.1537e-10 | -4.9999e-05 | 9.7254e-10 | 9.7254e-10 | 5.9e-09 | -5.9e-09 |  | 6.9396e-10 | 11.76%
-2e-05 | 2e-05 | -2e-05 | 4.5104e-10 | 4.5104e-10 | -1.9999e-05 | 8.1837e-10 | 8.1837e-10 | 4.34e-09 | -4.34e-09 |  | 6.3471e-10 | 14.62%
-1e-05 | 1e-05 | -9.9992e-06 | 8.1061e-10 | 8.1061e-10 | -9.999e-06 | 1.0139e-09 | 1.0139e-09 | 3.82e-09 | -3.82e-09 |  | 9.1225e-10 | 23.88%
-5e-06 | 5e-06 | -4.9992e-06 | 8.4144e-10 | 8.4144e-10 | -4.999e-06 | 9.9993e-10 | 9.9993e-10 | 3.56e-09 | -3.56e-09 |  | 9.2068e-10 | 25.86%
-2e-06 | 2e-06 | -1.9991e-06 | 8.604e-10 | 8.604e-10 | -1.9989e-06 | 1.1087e-09 | 1.1087e-09 | 3.404e-09 | -3.404e-09 |  | 9.8455e-10 | 28.92%
-1e-06 | 1e-06 | -9.9915e-07 | 8.4893e-10 | 8.4893e-10 | -9.9882e-07 | 1.1793e-09 | 1.1793e-09 | 3.352e-09 | -3.352e-09 |  | 1.0141e-09 | 30.25%
-5e-07 | 5e-07 | -4.9921e-07 | 7.9071e-10 | 7.9071e-10 | -4.9876e-07 | 1.2403e-09 | 1.2403e-09 | 3.326e-09 | -3.326e-09 |  | 1.0155e-09 | 30.53%
-2e-07 | 2e-07 | -1.9922e-07 | 7.8324e-10 | 7.8324e-10 | -1.9872e-07 | 1.2808e-09 | 1.2808e-09 | 3.3104e-09 | -3.3104e-09 |  | 1.032e-09 | 31.17%
-1e-07 | 1e-07 | -9.9181e-08 | 8.185e-10 | 8.185e-10 | -9.8918e-08 | 1.0818e-09 | 1.0818e-09 | 3.3052e-09 | -3.3052e-09 |  | 9.5017e-10 | 28.75%
-5e-08 | 5e-08 | -4.9139e-08 | 8.6118e-10 | 8.6118e-10 | -4.8823e-08 | 1.1768e-09 | 1.1768e-09 | 3.3026e-09 | -3.3026e-09 |  | 1.019e-09 | 30.85%
-2e-08 | 2e-08 | -1.9177e-08 | 8.2268e-10 | 8.2268e-10 | -1.9001e-08 | 9.9891e-10 | 9.9891e-10 | 3.301e-09 | -3.301e-09 |  | 9.108e-10 | 27.59%
-1e-08 | 1e-08 | -9.1593e-09 | 8.4074e-10 | 8.4074e-10 | -9.0407e-09 | 9.5929e-10 | 9.5929e-10 | 3.3005e-09 | -3.3005e-09 |  | 9.0002e-10 | 27.27%
-5e-09 | 5e-09 | -4.1832e-09 | 8.1678e-10 | 8.1678e-10 | -4.0322e-09 | 9.6779e-10 | 9.6779e-10 | 3.3003e-09 | -3.3003e-09 |  | 8.9229e-10 | 27.04%
-2e-09 | 2e-09 | -1.1923e-09 | 8.0774e-10 | 8.0774e-10 | -1.0453e-09 | 9.5473e-10 | 9.5473e-10 | 3.3001e-09 | -3.3001e-09 |  | 8.8123e-10 | 26.7%
-1e-09 | 1e-09 | -1.7977e-10 | 8.2023e-10 | 8.2023e-10 | -7.3963e-11 | 9.2604e-10 | 9.2604e-10 | 3.3001e-09 | -3.3001e-09 |  | 8.7313e-10 | 26.46%
-5e-10 | 5e-10 | 3.0518e-10 | 8.0518e-10 | 8.0518e-10 | 4.2457e-10 | 9.2457e-10 | 9.2457e-10 | 3.3e-09 | -3.3e-09 |  | 8.6488e-10 | 26.21%
-2e-10 | 2e-10 | 5.9701e-10 | 7.9701e-10 | 7.9701e-10 | 7.192e-10 | 9.192e-10 | 9.192e-10 | 3.3e-09 | -3.3e-09 |  | 8.581e-10 | 26%
-1e-10 | 1e-10 | 6.9213e-10 | 7.9213e-10 | 7.9213e-10 | 7.87e-10 | 8.87e-10 | 8.87e-10 | 3.3e-09 | -3.3e-09 |  | 8.3957e-10 | 25.44%
0 | 0 | 7.3944e-10 | 7.3944e-10 | 7.3944e-10 | 8.7107e-10 | 8.7107e-10 | 8.7107e-10 | 3.3e-09 | -3.3e-09 |  | 8.0525e-10 | 24.4%
1e-10 | 1e-10 | 8.5844e-10 | 7.5844e-10 | 7.5844e-10 | 9.8366e-10 | 8.8366e-10 | 8.8366e-10 | 3.3e-09 | -3.3e-09 |  | 8.2105e-10 | 24.88%
2e-10 | 2e-10 | 9.9088e-10 | 7.9088e-10 | 7.9088e-10 | 1.1011e-09 | 9.0112e-10 | 9.0112e-10 | 3.3e-09 | -3.3e-09 |  | 8.46e-10 | 25.64%
5e-10 | 5e-10 | 1.3122e-09 | 8.1223e-10 | 8.1223e-10 | 1.4248e-09 | 9.2484e-10 | 9.2484e-10 | 3.3e-09 | -3.3e-09 |  | 8.6853e-10 | 26.32%
1e-09 | 1e-09 | 1.8093e-09 | 8.0935e-10 | 8.0935e-10 | 1.9406e-09 | 9.4065e-10 | 9.4065e-10 | 3.3001e-09 | -3.3001e-09 |  | 8.75e-10 | 26.51%
2e-09 | 2e-09 | 2.8232e-09 | 8.2317e-10 | 8.2317e-10 | 2.9261e-09 | 9.2605e-10 | 9.2605e-10 | 3.3001e-09 | -3.3001e-09 |  | 8.7461e-10 | 26.5%
5e-09 | 5e-09 | 5.8439e-09 | 8.4386e-10 | 8.4386e-10 | 5.9181e-09 | 9.1809e-10 | 9.1809e-10 | 3.3003e-09 | -3.3003e-09 |  | 8.8098e-10 | 26.69%
1e-08 | 1e-08 | 1.0822e-08 | 8.2233e-10 | 8.2233e-10 | 1.1275e-08 | 1.2747e-09 | 1.2747e-09 | 3.3005e-09 | -3.3005e-09 |  | 1.0485e-09 | 31.77%
2e-08 | 2e-08 | 2.0839e-08 | 8.3855e-10 | 8.3855e-10 | 2.1237e-08 | 1.2373e-09 | 1.2373e-09 | 3.301e-09 | -3.301e-09 |  | 1.0379e-09 | 31.44%
5e-08 | 5e-08 | 5.0832e-08 | 8.3175e-10 | 8.3175e-10 | 5.118e-08 | 1.1803e-09 | 1.1803e-09 | 3.3026e-09 | -3.3026e-09 |  | 1.006e-09 | 30.46%
1e-07 | 1e-07 | 1.008e-07 | 7.9641e-10 | 7.9641e-10 | 1.0105e-07 | 1.0534e-09 | 1.0534e-09 | 3.3052e-09 | -3.3052e-09 |  | 9.2492e-10 | 27.98%
2e-07 | 2e-07 | 2.0078e-07 | 7.8312e-10 | 7.8312e-10 | 2.0101e-07 | 1.0119e-09 | 1.0119e-09 | 3.3104e-09 | -3.3104e-09 |  | 8.975e-10 | 27.11%
5e-07 | 5e-07 | 5.0079e-07 | 7.9117e-10 | 7.9117e-10 | 5.011e-07 | 1.1024e-09 | 1.1024e-09 | 3.326e-09 | -3.326e-09 |  | 9.4677e-10 | 28.47%
1e-06 | 1e-06 | 1.0009e-06 | 8.5439e-10 | 8.5439e-10 | 1.0011e-06 | 1.1108e-09 | 1.1108e-09 | 3.352e-09 | -3.352e-09 |  | 9.8258e-10 | 29.31%
2e-06 | 2e-06 | 2.0008e-06 | 8.4253e-10 | 8.4253e-10 | 2.0012e-06 | 1.2166e-09 | 1.2166e-09 | 3.404e-09 | -3.404e-09 |  | 1.0296e-09 | 30.25%
5e-06 | 5e-06 | 5.0008e-06 | 7.9768e-10 | 7.9768e-10 | 5.0012e-06 | 1.2324e-09 | 1.2324e-09 | 3.56e-09 | -3.56e-09 |  | 1.015e-09 | 28.51%
1e-05 | 1e-05 | 1.0001e-05 | 5.5989e-10 | 5.5989e-10 | 1.0001e-05 | 9.6551e-10 | 9.6551e-10 | 3.82e-09 | -3.82e-09 |  | 7.627e-10 | 19.97%
2e-05 | 2e-05 | 2.0001e-05 | 6.0369e-10 | 6.0369e-10 | 2.0001e-05 | 1.0231e-09 | 1.0231e-09 | 4.34e-09 | -4.34e-09 |  | 8.1337e-10 | 18.74%
5e-05 | 5e-05 | 5.0001e-05 | 6.1822e-10 | 6.1822e-10 | 5.0001e-05 | 9.7757e-10 | 9.7757e-10 | 5.9e-09 | -5.9e-09 |  | 7.9789e-10 | 13.52%
0.0001 | 0.0001 | 0.0001 | 6.936e-10 | 6.936e-10 | 0.0001 | 1.1475e-09 | 1.1475e-09 | 8.5e-09 | -8.5e-09 |  | 9.2055e-10 | 10.83%
0.0002 | 0.0002 | 0.00020001 | 5.1313e-09 | 5.1313e-09 | 0.00020001 | 8.3828e-09 | 8.3828e-09 | 3.04e-08 | -3.04e-08 |  | 6.757e-09 | 22.23%
0.0005 | 0.0005 | 0.0005 | 4.762e-09 | 4.762e-09 | 0.00050001 | 7.202e-09 | 7.202e-09 | 4.6e-08 | -4.6e-08 |  | 5.982e-09 | 13%
0.001 | 0.001 | 0.001 | 6.129e-09 | 6.129e-09 | 0.001 | 9.162e-09 | 9.162e-09 | 7.2e-08 | -7.2e-08 |  | 7.6455e-09 | 10.62%
0.002 | 0.002 | 0.002 | 3.1153e-08 | 3.1153e-08 | 0.0020001 | 6.713e-08 | 6.713e-08 | 3.68e-07 | -3.68e-07 |  | 4.9142e-08 | 13.35%
0.005 | 0.005 | 0.005 | 1.8923e-08 | 1.8923e-08 | 0.0050001 | 6.2537e-08 | 6.2537e-08 | 5.9e-07 | -5.9e-07 |  | 4.073e-08 | 6.903%
0.01 | 0.01 | 0.01 | 1.838e-08 | 1.838e-08 | 0.01 | 3.972e-08 | 3.972e-08 | 9.6e-07 | -9.6e-07 |  | 2.905e-08 | 3.026%
0.02 | 0.02 | 0.020001 | 7.6558e-07 | 7.6558e-07 | 0.020001 | 6.4771e-07 | 6.4771e-07 | 7.34e-06 | -7.34e-06 |  | 7.0665e-07 | 9.627%
0.05 | 0.05 | 0.050001 | 1.0333e-06 | 1.0333e-06 | 0.050001 | 1.0078e-06 | 1.0078e-06 | 1.34e-05 | -1.34e-05 |  | 1.0205e-06 | 7.616%
0.1 | 0.1 | 0.1 | 9.521e-07 | 9.521e-07 | 0.1 | 5.838e-07 | 5.838e-07 | 2.35e-05 | -2.35e-05 |  | 7.6795e-07 | 3.268%
[11 empty] | overall | 21.46%
average

## calibration results - 2015-12-1
run at 2015-12-12 17-46-25
channel | A/B
temperature | 34.6
line frequency | 49.9891
auto zero | TRUE
DCV
Channel A/B | Channel A/B | Channel A | Channel A | Channel A | Channel A 2nd run | Channel A 2nd run | Channel A 2nd run | Channel B | Channel B | Channel B | Channel B 2nd run | Channel B 2nd run | Channel B 2nd run | specified | -specified | average | average(Error)/
targetValue | abs(targetValue) | result | error (uV) | abs(Error) | result | error (uV) | abs(Error) | result | error (uV) | abs(Error) | result | error (uV) | abs(Error) | error (uV) | error (uV) | error | specified error
-10 | 10 | -10 | -8 | 8 | -10 | 14 | 14 | -10 | -26 | 26 | -10 | -25 | 25 | 560.4878 | -560.4878
-5 | 5 | -5 | 15 | 15 | -5 | 24 | 24 | -5 | -6 | 6 | -5 | -1 | 1 | 370.2439 | -370.2439
-2 | 2 | -2 | 15 | 15 | -2 | 20 | 20 | -2 | -12 | 12 | -2 | -12 | 12 | 256.0976 | -256.0976
-1 | 1 | -0.99999 | 5 | 5 | -0.99999 | 7 | 7 | -1 | 3 | 3 | -1 | 2 | 2 | 73.2 | -73.2
-0.5 | 0.5 | -0.49999 | 12 | 12 | -0.49999 | 11 | 11 | -0.49999 | 8 | 8 | -0.49999 | 8 | 8 | 51.8 | -51.8
-0.2 | 0.2 | -0.19999 | 11 | 11 | -0.19999 | 10 | 10 | -0.19999 | 9 | 9 | -0.19999 | 8 | 8 | 39.32 | -39.32
-0.1 | 0.1 | -0.099991 | 9 | 9 | -0.099991 | 9 | 9 | -0.09999 | 10 | 10 | -0.099991 | 9 | 9 | 35.16 | -35.16
-0.05 | 0.05 | -0.049989 | 11 | 11 | -0.04999 | 10 | 10 | -0.049991 | 9 | 9 | -0.049992 | 8 | 8 | 33.08 | -33.08
-0.02 | 0.02 | -0.019992 | 8 | 8 | -0.01999 | 10 | 10 | -0.019992 | 8 | 8 | -0.019992 | 8 | 8 | 31.832 | -31.832
-0.01 | 0.01 | -0.0099903 | 10 | 10 | -0.0099903 | 10 | 10 | -0.0099904 | 10 | 10 | -0.0099912 | 9 | 9 | 31.416 | -31.416
-0.005 | 0.005 | -0.0049913 | 9 | 9 | -0.0049902 | 10 | 10 | -0.0049934 | 7 | 7 | -0.0049946 | 5 | 5 | 31.208 | -31.208
-0.002 | 0.002 | -0.0019927 | 7 | 7 | -0.0019909 | 9 | 9 | -0.0019901 | 10 | 10 | -0.0019925 | 8 | 8 | 31.0832 | -31.0832
-0.001 | 0.001 | -0.00099063 | 9 | 9 | -0.00099051 | 9 | 9 | -0.00099014 | 10 | 10 | -0.00099136 | 9 | 9 | 31.0416 | -31.0416
-0.0005 | 0.0005 | -0.00049331 | 7 | 7 | -0.0004926 | 7 | 7 | -0.00049174 | 8 | 8 | -0.00049272 | 7 | 7 | 31.0208 | -31.0208
-0.0002 | 0.0002 | -0.0001922 | 8 | 8 | -0.00019345 | 7 | 7 | -0.00019306 | 7 | 7 | -0.0001942 | 6 | 6 | 31.0083 | -31.0083
-0.0001 | 0.0001 | -9.2066e-05 | 8 | 8 | -9.3761e-05 | 6 | 6 | -9.3351e-05 | 7 | 7 | -9.463e-05 | 5 | 5 | 31.0042 | -31.0042
-5e-05 | 5e-05 | -4.1803e-05 | 8 | 8 | -4.0699e-05 | 9 | 9 | -4.3638e-05 | 6 | 6 | -4.4746e-05 | 5 | 5 | 31.0021 | -31.0021
-2e-05 | 2e-05 | -1.1859e-05 | 8 | 8 | -1.0532e-05 | 9 | 9 | -1.3503e-05 | 6 | 6 | -1.4258e-05 | 6 | 6 | 31.0008 | -31.0008
-1e-05 | 1e-05 | -2.3715e-06 | 8 | 8 | -8.754e-07 | 9 | 9 | -3.8814e-06 | 6 | 6 | -4.8128e-06 | 5 | 5 | 31.0004 | -31.0004
-5e-06 | 5e-06 | 2.663e-06 | 8 | 8 | 3.8846e-06 | 9 | 9 | 1.1609e-06 | 6 | 6 | 1.9986e-07 | 5 | 5 | 31.0002 | -31.0002
-2e-06 | 2e-06 | 5.2196e-06 | 7 | 7 | 6.284e-06 | 8 | 8 | 3.7753e-06 | 6 | 6 | 3.1141e-06 | 5 | 5 | 31.0001 | -31.0001
-1e-06 | 1e-06 | 5.7081e-06 | 7 | 7 | 6.486e-06 | 7 | 7 | 4.5977e-06 | 6 | 6 | 4.0083e-06 | 5 | 5 | 31 | -31
0 | 0 | 6.1536e-06 | 6 | 6 | 6.2313e-06 | 6 | 6 | 5.534e-06 | 6 | 6 | 4.9878e-06 | 5 | 5 | 31 | -31
1e-06 | 1e-06 | 7.2031e-06 | 6 | 6 | 8.2929e-06 | 7 | 7 | 6.4423e-06 | 5 | 5 | 5.9065e-06 | 5 | 5 | 31 | -31 | 5.75 | 18.55%
2e-06 | 2e-06 | 8.5623e-06 | 7 | 7 | 9.855e-06 | 8 | 8 | 7.877e-06 | 6 | 6 | 7.0102e-06 | 5 | 5 | 31.0001 | -31.0001 | 6.5 | 20.97%
5e-06 | 5e-06 | 1.2229e-05 | 7 | 7 | 1.3666e-05 | 9 | 9 | 1.0786e-05 | 6 | 6 | 1.0102e-05 | 5 | 5 | 31.0002 | -31.0002 | 6.75 | 21.77%
1e-05 | 1e-05 | 1.7807e-05 | 8 | 8 | 1.8559e-05 | 9 | 9 | 1.5842e-05 | 6 | 6 | 1.5335e-05 | 5 | 5 | 31.0004 | -31.0004 | 7 | 22.58%
2e-05 | 2e-05 | 2.7728e-05 | 8 | 8 | 2.8811e-05 | 9 | 9 | 2.6342e-05 | 6 | 6 | 2.5191e-05 | 5 | 5 | 31.0008 | -31.0008 | 7 | 22.58%
5e-05 | 5e-05 | 5.838e-05 | 8 | 8 | 5.9386e-05 | 9 | 9 | 5.6143e-05 | 6 | 6 | 5.5451e-05 | 5 | 5 | 31.0021 | -31.0021 | 7 | 22.58%
0.0001 | 0.0001 | 0.00010867 | 9 | 9 | 0.00010917 | 9 | 9 | 0.00010592 | 6 | 6 | 0.00010511 | 5 | 5 | 31.0042 | -31.0042 | 7.25 | 23.38%
0.0002 | 0.0002 | 0.00020897 | 9 | 9 | 0.00020896 | 9 | 9 | 0.00020561 | 6 | 6 | 0.00020478 | 5 | 5 | 31.0083 | -31.0083 | 7.25 | 23.38%
0.0005 | 0.0005 | 0.00051004 | 10 | 10 | 0.00050892 | 9 | 9 | 0.00050892 | 9 | 9 | 0.00050806 | 8 | 8 | 31.0208 | -31.0208 | 9 | 29.01%
0.001 | 0.001 | 0.0010087 | 9 | 9 | 0.0010087 | 9 | 9 | 0.0010083 | 8 | 8 | 0.0010071 | 7 | 7 | 31.0416 | -31.0416 | 8.25 | 26.58%
0.002 | 0.002 | 0.0020064 | 6 | 6 | 0.0020075 | 7 | 7 | 0.0020088 | 9 | 9 | 0.0020078 | 8 | 8 | 31.0832 | -31.0832 | 7.5 | 24.13%
0.005 | 0.005 | 0.0050076 | 8 | 8 | 0.005007 | 7 | 7 | 0.0050075 | 7 | 7 | 0.0050065 | 7 | 7 | 31.208 | -31.208 | 7.25 | 23.23%
0.01 | 0.01 | 0.010007 | 7 | 7 | 0.010008 | 8 | 8 | 0.010009 | 9 | 9 | 0.010007 | 7 | 7 | 31.416 | -31.416 | 7.75 | 24.67%
0.02 | 0.02 | 0.020006 | 6 | 6 | 0.020008 | 8 | 8 | 0.02001 | 10 | 10 | 0.020008 | 8 | 8 | 31.832 | -31.832 | 8 | 25.13%
0.05 | 0.05 | 0.050008 | 8 | 8 | 0.050008 | 8 | 8 | 0.050009 | 9 | 9 | 0.050009 | 9 | 9 | 33.08 | -33.08 | 8.5 | 25.7%
0.1 | 0.1 | 0.10001 | 10 | 10 | 0.10001 | 9 | 9 | 0.10001 | 8 | 8 | 0.10001 | 7 | 7 | 35.16 | -35.16 | 8.5 | 24.18%
0.2 | 0.2 | 0.20001 | 8 | 8 | 0.20001 | 8 | 8 | 0.20001 | 9 | 9 | 0.20001 | 8 | 8 | 39.32 | -39.32 | 8.25 | 20.98%
0.5 | 0.5 | 0.50001 | 6 | 6 | 0.50001 | 6 | 6 | 0.50001 | 8 | 8 | 0.50001 | 6 | 6 | 51.8 | -51.8 | 6.5 | 12.55%
1 | 1 | 1 | 4 | 4 | 1 | 5 | 5 | 1 | 4 | 4 | 1 | 3 | 3 | 73.2 | -73.2 | 4 | 5.464%
2 | 2 | 2 | 18 | 18 | 2 | 23 | 23 | 2 | 9 | 9 | 2 | 12 | 12 | 256.0976 | -256.0976 | 15.5 | 6.052%
5 | 5 | 5 | 16 | 16 | 5 | 23 | 23 | 5 | -4 | 4 | 5 | -3 | 3 | 370.2439 | -370.2439 | 11.5 | 3.106%
10 | 10 | 10 | 10 | 10 | 10 | 23 | 23 | 10 | -15 | 15 | 10 | -16 | 16 | 560.4878 | -560.4878 | 16 | 2.855%
[16 empty] | overall | 19.52%
average
DCI
Channel A/B | Channel A/B | Channel A | Channel A | Channel A |  | Channel B | Channel B | Channel B |  | specified | -specified |  | average | average(Error)/
targetValue | abs(targetValue) | result | error (A) | abs(Error) |  | result | error (A) | abs(Error) |  | error (A) | error (A) |  | error | specified error
-0.1 | 0.1 | -0.1 | -7.41e-08 | 7.41e-08 |  | -0.099999 | 1.0591e-06 | 1.0591e-06 |  | 2.35e-05 | -2.35e-05 |  | 5.6662e-07 | 2.411%
-0.05 | 0.05 | -0.05 | 1.2842e-07 | 1.2842e-07 |  | -0.05 | 3.3372e-07 | 3.3372e-07 |  | 1.34e-05 | -1.34e-05 |  | 2.3107e-07 | 1.724%
-0.02 | 0.02 | -0.02 | -5.939e-08 | 5.939e-08 |  | -0.02 | 2.9449e-07 | 2.9449e-07 |  | 7.34e-06 | -7.34e-06 |  | 1.7694e-07 | 2.411%
-0.01 | 0.01 | -0.01 | -9.37e-09 | 9.37e-09 |  | -0.01 | 2.8875e-08 | 2.8875e-08 |  | 9.6e-07 | -9.6e-07 |  | 1.9122e-08 | 1.992%
-0.005 | 0.005 | -0.005 | -1.9604e-08 | 1.9604e-08 |  | -0.005 | 1.3694e-08 | 1.3694e-08 |  | 5.9e-07 | -5.9e-07 |  | 1.6649e-08 | 2.822%
-0.002 | 0.002 | -0.002 | -1.8718e-08 | 1.8718e-08 |  | -0.002 | 3.2834e-08 | 3.2834e-08 |  | 3.68e-07 | -3.68e-07 |  | 2.5776e-08 | 7.004%
-0.001 | 0.001 | -0.001 | 2.096e-10 | 2.096e-10 |  | -0.001 | 2.674e-09 | 2.674e-09 |  | 7.2e-08 | -7.2e-08 |  | 1.4418e-09 | 2.002%
-0.0005 | 0.0005 | -0.0005 | 1.3519e-09 | 1.3519e-09 |  | -0.00049999 | 5.0708e-09 | 5.0708e-09 |  | 4.6e-08 | -4.6e-08 |  | 3.2114e-09 | 6.981%
-0.0002 | 0.0002 | -0.0002 | -1.2064e-09 | 1.2064e-09 |  | -0.0002 | 2.9129e-09 | 2.9129e-09 |  | 3.04e-08 | -3.04e-08 |  | 2.0597e-09 | 6.775%
-0.0001 | 0.0001 | -0.0001 | -4.64e-11 | 4.64e-11 |  | -0.0001 | 3.1669e-10 | 3.1669e-10 |  | 8.5e-09 | -8.5e-09 |  | 1.8155e-10 | 2.136%
-5e-05 | 5e-05 | -5e-05 | 6.822e-11 | 6.822e-11 |  | -5e-05 | 4.5832e-10 | 4.5832e-10 |  | 5.9e-09 | -5.9e-09 |  | 2.6327e-10 | 4.462%
-2e-05 | 2e-05 | -2e-05 | -7.891e-11 | 7.891e-11 |  | -2e-05 | 4.5899e-10 | 4.5899e-10 |  | 4.34e-09 | -4.34e-09 |  | 2.6895e-10 | 6.197%
-1e-05 | 1e-05 | -1e-05 | -5.382e-11 | 5.382e-11 |  | -9.9997e-06 | 2.9691e-10 | 2.9691e-10 |  | 3.82e-09 | -3.82e-09 |  | 1.7537e-10 | 4.591%
-5e-06 | 5e-06 | -5e-06 | 2.2461e-11 | 2.2461e-11 |  | -4.9997e-06 | 2.8276e-10 | 2.8276e-10 |  | 3.56e-09 | -3.56e-09 |  | 1.5261e-10 | 4.287%
-2e-06 | 2e-06 | -2e-06 | 3.7955e-11 | 3.7955e-11 |  | -1.9997e-06 | 3.3422e-10 | 3.3422e-10 |  | 3.404e-09 | -3.404e-09 |  | 1.8609e-10 | 5.467%
-1e-06 | 1e-06 | -1.0002e-06 | -2.0799e-10 | 2.0799e-10 |  | -9.9955e-07 | 4.4946e-10 | 4.4946e-10 |  | 3.352e-09 | -3.352e-09 |  | 3.2873e-10 | 9.807%
-5e-07 | 5e-07 | -5.0006e-07 | -6.206e-11 | 6.206e-11 |  | -4.995e-07 | 5.0462e-10 | 5.0462e-10 |  | 3.326e-09 | -3.326e-09 |  | 2.8334e-10 | 8.519%
-2e-07 | 2e-07 | -2.0023e-07 | -2.316e-10 | 2.316e-10 |  | -1.998e-07 | 1.9894e-10 | 1.9894e-10 |  | 3.3104e-09 | -3.3104e-09 |  | 2.1527e-10 | 6.503%
-1e-07 | 1e-07 | -1.0009e-07 | -8.8387e-11 | 8.8387e-11 |  | -9.9739e-08 | 2.6089e-10 | 2.6089e-10 |  | 3.3052e-09 | -3.3052e-09 |  | 1.7464e-10 | 5.284%
-5e-08 | 5e-08 | -5.0006e-08 | -6.2392e-12 | 6.2392e-12 |  | -4.9629e-08 | 3.7113e-10 | 3.7113e-10 |  | 3.3026e-09 | -3.3026e-09 |  | 1.8868e-10 | 5.713%
-2e-08 | 2e-08 | -2.0185e-08 | -1.8472e-10 | 1.8472e-10 |  | -1.9753e-08 | 2.4702e-10 | 2.4702e-10 |  | 3.301e-09 | -3.301e-09 |  | 2.1587e-10 | 6.539%
-1e-08 | 1e-08 | -1.0194e-08 | -1.9404e-10 | 1.9404e-10 |  | -9.7727e-09 | 2.2732e-10 | 2.2732e-10 |  | 3.3005e-09 | -3.3005e-09 |  | 2.1068e-10 | 6.383%
-5e-09 | 5e-09 | -5.171e-09 | -1.7102e-10 | 1.7102e-10 |  | -4.8033e-09 | 1.9668e-10 | 1.9668e-10 |  | 3.3003e-09 | -3.3003e-09 |  | 1.8385e-10 | 5.571%
-2e-09 | 2e-09 | -2.1633e-09 | -1.6333e-10 | 1.6333e-10 |  | -1.8012e-09 | 1.9881e-10 | 1.9881e-10 |  | 3.3001e-09 | -3.3001e-09 |  | 1.8107e-10 | 5.487%
-1e-09 | 1e-09 | -1.1855e-09 | -1.8546e-10 | 1.8546e-10 |  | -7.8507e-10 | 2.1493e-10 | 2.1493e-10 |  | 3.3001e-09 | -3.3001e-09 |  | 2.002e-10 | 6.066%
-5e-10 | 5e-10 | -6.6938e-10 | -1.6938e-10 | 1.6938e-10 |  | -2.8771e-10 | 2.1229e-10 | 2.1229e-10 |  | 3.3e-09 | -3.3e-09 |  | 1.9083e-10 | 5.783%
-2e-10 | 2e-10 | -3.7335e-10 | -1.7335e-10 | 1.7335e-10 |  | -1.2572e-12 | 1.9874e-10 | 1.9874e-10 |  | 3.3e-09 | -3.3e-09 |  | 1.8605e-10 | 5.638%
-1e-10 | 1e-10 | -2.9393e-10 | -1.9393e-10 | 1.9393e-10 |  | 1.06e-10 | 2.06e-10 | 2.06e-10 |  | 3.3e-09 | -3.3e-09 |  | 1.9996e-10 | 6.059%
0 | 0 | -1.8866e-10 | -1.8866e-10 | 1.8866e-10 |  | 2.2059e-10 | 2.2059e-10 | 2.2059e-10 |  | 3.3e-09 | -3.3e-09 |  | 2.0463e-10 | 6.201%
1e-10 | 1e-10 | -7.8807e-11 | -1.7881e-10 | 1.7881e-10 |  | 2.8707e-10 | 1.8707e-10 | 1.8707e-10 |  | 3.3e-09 | -3.3e-09 |  | 1.8294e-10 | 5.544%
2e-10 | 2e-10 | 1.2397e-11 | -1.876e-10 | 1.876e-10 |  | 4.1801e-10 | 2.1801e-10 | 2.1801e-10 |  | 3.3e-09 | -3.3e-09 |  | 2.0281e-10 | 6.146%
5e-10 | 5e-10 | 3.1851e-10 | -1.8149e-10 | 1.8149e-10 |  | 6.6646e-10 | 1.6646e-10 | 1.6646e-10 |  | 3.3e-09 | -3.3e-09 |  | 1.7398e-10 | 5.272%
1e-09 | 1e-09 | 8.2467e-10 | -1.7533e-10 | 1.7533e-10 |  | 1.1955e-09 | 1.9552e-10 | 1.9552e-10 |  | 3.3001e-09 | -3.3001e-09 |  | 1.8543e-10 | 5.619%
2e-09 | 2e-09 | 1.7974e-09 | -2.0257e-10 | 2.0257e-10 |  | 2.2073e-09 | 2.0734e-10 | 2.0734e-10 |  | 3.3001e-09 | -3.3001e-09 |  | 2.0496e-10 | 6.211%
5e-09 | 5e-09 | 4.8061e-09 | -1.939e-10 | 1.939e-10 |  | 5.2021e-09 | 2.0213e-10 | 2.0213e-10 |  | 3.3003e-09 | -3.3003e-09 |  | 1.9802e-10 | 6%
1e-08 | 1e-08 | 9.8461e-09 | -1.5393e-10 | 1.5393e-10 |  | 1.0529e-08 | 5.2929e-10 | 5.2929e-10 |  | 3.3005e-09 | -3.3005e-09 |  | 3.4161e-10 | 10.35%
2e-08 | 2e-08 | 1.9842e-08 | -1.5789e-10 | 1.5789e-10 |  | 2.048e-08 | 4.801e-10 | 4.801e-10 |  | 3.301e-09 | -3.301e-09 |  | 3.19e-10 | 9.664%
5e-08 | 5e-08 | 5.0009e-08 | 8.7791e-12 | 8.7791e-12 |  | 5.0375e-08 | 3.7527e-10 | 3.7527e-10 |  | 3.3026e-09 | -3.3026e-09 |  | 1.9202e-10 | 5.814%
1e-07 | 1e-07 | 9.9953e-08 | -4.6987e-11 | 4.6987e-11 |  | 1.0026e-07 | 2.593e-10 | 2.593e-10 |  | 3.3052e-09 | -3.3052e-09 |  | 1.5314e-10 | 4.633%
2e-07 | 2e-07 | 1.9981e-07 | -1.9281e-10 | 1.9281e-10 |  | 2.0027e-07 | 2.6907e-10 | 2.6907e-10 |  | 3.3104e-09 | -3.3104e-09 |  | 2.3094e-10 | 6.976%
5e-07 | 5e-07 | 4.9994e-07 | -5.7659e-11 | 5.7659e-11 |  | 5.0033e-07 | 3.2513e-10 | 3.2513e-10 |  | 3.326e-09 | -3.326e-09 |  | 1.914e-10 | 5.755%
1e-06 | 1e-06 | 1e-06 | 3.1426e-11 | 3.1426e-11 |  | 1.0004e-06 | 3.959e-10 | 3.959e-10 |  | 3.352e-09 | -3.352e-09 |  | 2.1366e-10 | 6.374%
2e-06 | 2e-06 | 2e-06 | 2.1608e-11 | 2.1608e-11 |  | 2.0005e-06 | 5.0645e-10 | 5.0645e-10 |  | 3.404e-09 | -3.404e-09 |  | 2.6403e-10 | 7.756%
5e-06 | 5e-06 | 5e-06 | 1.7172e-11 | 1.7172e-11 |  | 5.0002e-06 | 2.1839e-10 | 2.1839e-10 |  | 3.56e-09 | -3.56e-09 |  | 1.1778e-10 | 3.308%
1e-05 | 1e-05 | 1e-05 | 6.247e-11 | 6.247e-11 |  | 1.0001e-05 | 5.7097e-10 | 5.7097e-10 |  | 3.82e-09 | -3.82e-09 |  | 3.1672e-10 | 8.291%
2e-05 | 2e-05 | 2e-05 | -2.936e-11 | 2.936e-11 |  | 2.0001e-05 | 5.3003e-10 | 5.3003e-10 |  | 4.34e-09 | -4.34e-09 |  | 2.7969e-10 | 6.445%
5e-05 | 5e-05 | 5e-05 | 1.4244e-10 | 1.4244e-10 |  | 5.0001e-05 | 5.4794e-10 | 5.4794e-10 |  | 5.9e-09 | -5.9e-09 |  | 3.4519e-10 | 5.851%
0.0001 | 0.0001 | 0.0001 | 1.198e-10 | 1.198e-10 |  | 0.0001 | 3.135e-10 | 3.135e-10 |  | 8.5e-09 | -8.5e-09 |  | 2.1665e-10 | 2.549%
0.0002 | 0.0002 | 0.0002 | -8.875e-10 | 8.875e-10 |  | 0.0002 | 4.8003e-09 | 4.8003e-09 |  | 3.04e-08 | -3.04e-08 |  | 2.8439e-09 | 9.355%
0.0005 | 0.0005 | 0.0005 | -7.21e-11 | 7.21e-11 |  | 0.0005 | 3.5442e-09 | 3.5442e-09 |  | 4.6e-08 | -4.6e-08 |  | 1.8082e-09 | 3.931%
0.001 | 0.001 | 0.001 | 2.83e-10 | 2.83e-10 |  | 0.001 | 1.66e-09 | 1.66e-09 |  | 7.2e-08 | -7.2e-08 |  | 9.715e-10 | 1.349%
0.002 | 0.002 | 0.002 | -3.7166e-08 | 3.7166e-08 |  | 0.002 | 2.853e-08 | 2.853e-08 |  | 3.68e-07 | -3.68e-07 |  | 3.2848e-08 | 8.926%
0.005 | 0.005 | 0.005 | -2.853e-08 | 2.853e-08 |  | 0.005 | 4.5151e-08 | 4.5151e-08 |  | 5.9e-07 | -5.9e-07 |  | 3.684e-08 | 6.244%
0.01 | 0.01 | 0.01 | -2.237e-09 | 2.237e-09 |  | 0.01 | 3.035e-08 | 3.035e-08 |  | 9.6e-07 | -9.6e-07 |  | 1.6293e-08 | 1.697%
0.02 | 0.02 | 0.02 | 1.9176e-07 | 1.9176e-07 |  | 0.02 | 1.0126e-07 | 1.0126e-07 |  | 7.34e-06 | -7.34e-06 |  | 1.4651e-07 | 1.996%
0.05 | 0.05 | 0.05 | 4.4292e-07 | 4.4292e-07 |  | 0.05 | 4.7939e-07 | 4.7939e-07 |  | 1.34e-05 | -1.34e-05 |  | 4.6115e-07 | 3.441%
0.1 | 0.1 | 0.1 | 3.996e-07 | 3.996e-07 |  | 0.1 | -2.771e-07 | 2.771e-07 |  | 2.35e-05 | -2.35e-05 |  | 3.3835e-07 | 1.44%
[13 empty] | overall | 5.652%
average