## calibration results - 2015-12-1
run at 2015-12-12 17-46-25
channel | A/B
temperature | 34.6
line frequency | 49.9891
auto zero | TRUE
Channel A/B | Channel A/B | Channel A | Channel A | Channel A | Channel A 2nd run | Channel A 2nd run | Channel A 2nd run | Channel B | Channel B | Channel B | Channel B 2nd run | Channel B 2nd run | Channel B 2nd run | specified | -specified | average | average(Error)/
targetValue | abs(targetValue) | result | error (uV) | abs(Error) | result | error (uV) | abs(Error) | result | error (uV) | abs(Error) | result | error (uV) | abs(Error) | error (uV) | error (uV) | error | specified error
-10 | 10 | -10 | -8 | 8 | -10 | 14 | 14 | -10 | -26 | 26 | -10 | -25 | 25 | 560.4878 | -560.4878
-5 | 5 | -5 | 15 | 15 | -5 | 24 | 24 | -5 | -6 | 6 | -5 | -1 | 1 | 370.2439 | -370.2439
-2 | 2 | -2 | 15 | 15 | -2 | 20 | 20 | -2 | -12 | 12 | -2 | -12 | 12 | 256.0976 | -256.0976
-1 | 1 | -0.99999 | 5 | 5 | -0.99999 | 7 | 7 | -1 | 3 | 3 | -1 | 2 | 2 | 73.2 | -73.2
-0.5 | 0.5 | -0.49999 | 12 | 12 | -0.49999 | 11 | 11 | -0.49999 | 8 | 8 | -0.49999 | 8 | 8 | 51.8 | -51.8
-0.2 | 0.2 | -0.19999 | 11 | 11 | -0.19999 | 10 | 10 | -0.19999 | 9 | 9 | -0.19999 | 8 | 8 | 39.32 | -39.32
-0.1 | 0.1 | -0.099991 | 9 | 9 | -0.099991 | 9 | 9 | -0.09999 | 10 | 10 | -0.099991 | 9 | 9 | 35.16 | -35.16
-0.05 | 0.05 | -0.049989 | 11 | 11 | -0.04999 | 10 | 10 | -0.049991 | 9 | 9 | -0.049992 | 8 | 8 | 33.08 | -33.08
-0.02 | 0.02 | -0.019992 | 8 | 8 | -0.01999 | 10 | 10 | -0.019992 | 8 | 8 | -0.019992 | 8 | 8 | 31.832 | -31.832
-0.01 | 0.01 | -0.0099903 | 10 | 10 | -0.0099903 | 10 | 10 | -0.0099904 | 10 | 10 | -0.0099912 | 9 | 9 | 31.416 | -31.416
-0.005 | 0.005 | -0.0049913 | 9 | 9 | -0.0049902 | 10 | 10 | -0.0049934 | 7 | 7 | -0.0049946 | 5 | 5 | 31.208 | -31.208
-0.002 | 0.002 | -0.0019927 | 7 | 7 | -0.0019909 | 9 | 9 | -0.0019901 | 10 | 10 | -0.0019925 | 8 | 8 | 31.0832 | -31.0832
-0.001 | 0.001 | -0.00099063 | 9 | 9 | -0.00099051 | 9 | 9 | -0.00099014 | 10 | 10 | -0.00099136 | 9 | 9 | 31.0416 | -31.0416
-0.0005 | 0.0005 | -0.00049331 | 7 | 7 | -0.0004926 | 7 | 7 | -0.00049174 | 8 | 8 | -0.00049272 | 7 | 7 | 31.0208 | -31.0208
-0.0002 | 0.0002 | -0.0001922 | 8 | 8 | -0.00019345 | 7 | 7 | -0.00019306 | 7 | 7 | -0.0001942 | 6 | 6 | 31.0083 | -31.0083
-0.0001 | 0.0001 | -9.2066e-05 | 8 | 8 | -9.3761e-05 | 6 | 6 | -9.3351e-05 | 7 | 7 | -9.463e-05 | 5 | 5 | 31.0042 | -31.0042
-5e-05 | 5e-05 | -4.1803e-05 | 8 | 8 | -4.0699e-05 | 9 | 9 | -4.3638e-05 | 6 | 6 | -4.4746e-05 | 5 | 5 | 31.0021 | -31.0021
-2e-05 | 2e-05 | -1.1859e-05 | 8 | 8 | -1.0532e-05 | 9 | 9 | -1.3503e-05 | 6 | 6 | -1.4258e-05 | 6 | 6 | 31.0008 | -31.0008
-1e-05 | 1e-05 | -2.3715e-06 | 8 | 8 | -8.754e-07 | 9 | 9 | -3.8814e-06 | 6 | 6 | -4.8128e-06 | 5 | 5 | 31.0004 | -31.0004
-5e-06 | 5e-06 | 2.663e-06 | 8 | 8 | 3.8846e-06 | 9 | 9 | 1.1609e-06 | 6 | 6 | 1.9986e-07 | 5 | 5 | 31.0002 | -31.0002
-2e-06 | 2e-06 | 5.2196e-06 | 7 | 7 | 6.284e-06 | 8 | 8 | 3.7753e-06 | 6 | 6 | 3.1141e-06 | 5 | 5 | 31.0001 | -31.0001
-1e-06 | 1e-06 | 5.7081e-06 | 7 | 7 | 6.486e-06 | 7 | 7 | 4.5977e-06 | 6 | 6 | 4.0083e-06 | 5 | 5 | 31 | -31
0 | 0 | 6.1536e-06 | 6 | 6 | 6.2313e-06 | 6 | 6 | 5.534e-06 | 6 | 6 | 4.9878e-06 | 5 | 5 | 31 | -31
1e-06 | 1e-06 | 7.2031e-06 | 6 | 6 | 8.2929e-06 | 7 | 7 | 6.4423e-06 | 5 | 5 | 5.9065e-06 | 5 | 5 | 31 | -31 | 5.75 | 18.55%
2e-06 | 2e-06 | 8.5623e-06 | 7 | 7 | 9.855e-06 | 8 | 8 | 7.877e-06 | 6 | 6 | 7.0102e-06 | 5 | 5 | 31.0001 | -31.0001 | 6.5 | 20.97%
5e-06 | 5e-06 | 1.2229e-05 | 7 | 7 | 1.3666e-05 | 9 | 9 | 1.0786e-05 | 6 | 6 | 1.0102e-05 | 5 | 5 | 31.0002 | -31.0002 | 6.75 | 21.77%
1e-05 | 1e-05 | 1.7807e-05 | 8 | 8 | 1.8559e-05 | 9 | 9 | 1.5842e-05 | 6 | 6 | 1.5335e-05 | 5 | 5 | 31.0004 | -31.0004 | 7 | 22.58%
2e-05 | 2e-05 | 2.7728e-05 | 8 | 8 | 2.8811e-05 | 9 | 9 | 2.6342e-05 | 6 | 6 | 2.5191e-05 | 5 | 5 | 31.0008 | -31.0008 | 7 | 22.58%
5e-05 | 5e-05 | 5.838e-05 | 8 | 8 | 5.9386e-05 | 9 | 9 | 5.6143e-05 | 6 | 6 | 5.5451e-05 | 5 | 5 | 31.0021 | -31.0021 | 7 | 22.58%
0.0001 | 0.0001 | 0.00010867 | 9 | 9 | 0.00010917 | 9 | 9 | 0.00010592 | 6 | 6 | 0.00010511 | 5 | 5 | 31.0042 | -31.0042 | 7.25 | 23.38%
0.0002 | 0.0002 | 0.00020897 | 9 | 9 | 0.00020896 | 9 | 9 | 0.00020561 | 6 | 6 | 0.00020478 | 5 | 5 | 31.0083 | -31.0083 | 7.25 | 23.38%
0.0005 | 0.0005 | 0.00051004 | 10 | 10 | 0.00050892 | 9 | 9 | 0.00050892 | 9 | 9 | 0.00050806 | 8 | 8 | 31.0208 | -31.0208 | 9 | 29.01%
0.001 | 0.001 | 0.0010087 | 9 | 9 | 0.0010087 | 9 | 9 | 0.0010083 | 8 | 8 | 0.0010071 | 7 | 7 | 31.0416 | -31.0416 | 8.25 | 26.58%
0.002 | 0.002 | 0.0020064 | 6 | 6 | 0.0020075 | 7 | 7 | 0.0020088 | 9 | 9 | 0.0020078 | 8 | 8 | 31.0832 | -31.0832 | 7.5 | 24.13%
0.005 | 0.005 | 0.0050076 | 8 | 8 | 0.005007 | 7 | 7 | 0.0050075 | 7 | 7 | 0.0050065 | 7 | 7 | 31.208 | -31.208 | 7.25 | 23.23%
0.01 | 0.01 | 0.010007 | 7 | 7 | 0.010008 | 8 | 8 | 0.010009 | 9 | 9 | 0.010007 | 7 | 7 | 31.416 | -31.416 | 7.75 | 24.67%
0.02 | 0.02 | 0.020006 | 6 | 6 | 0.020008 | 8 | 8 | 0.02001 | 10 | 10 | 0.020008 | 8 | 8 | 31.832 | -31.832 | 8 | 25.13%
0.05 | 0.05 | 0.050008 | 8 | 8 | 0.050008 | 8 | 8 | 0.050009 | 9 | 9 | 0.050009 | 9 | 9 | 33.08 | -33.08 | 8.5 | 25.7%
0.1 | 0.1 | 0.10001 | 10 | 10 | 0.10001 | 9 | 9 | 0.10001 | 8 | 8 | 0.10001 | 7 | 7 | 35.16 | -35.16 | 8.5 | 24.18%
0.2 | 0.2 | 0.20001 | 8 | 8 | 0.20001 | 8 | 8 | 0.20001 | 9 | 9 | 0.20001 | 8 | 8 | 39.32 | -39.32 | 8.25 | 20.98%
0.5 | 0.5 | 0.50001 | 6 | 6 | 0.50001 | 6 | 6 | 0.50001 | 8 | 8 | 0.50001 | 6 | 6 | 51.8 | -51.8 | 6.5 | 12.55%
1 | 1 | 1 | 4 | 4 | 1 | 5 | 5 | 1 | 4 | 4 | 1 | 3 | 3 | 73.2 | -73.2 | 4 | 5.464%
2 | 2 | 2 | 18 | 18 | 2 | 23 | 23 | 2 | 9 | 9 | 2 | 12 | 12 | 256.0976 | -256.0976 | 15.5 | 6.052%
5 | 5 | 5 | 16 | 16 | 5 | 23 | 23 | 5 | -4 | 4 | 5 | -3 | 3 | 370.2439 | -370.2439 | 11.5 | 3.106%
10 | 10 | 10 | 10 | 10 | 10 | 23 | 23 | 10 | -15 | 15 | 10 | -16 | 16 | 560.4878 | -560.4878 | 16 | 2.855%
[16 empty] | overall | 19.52%
average
Channel A/B | Channel A/B | Channel A | Channel A | Channel A |  | Channel B | Channel B | Channel B |  | specified | -specified |  | average | average(Error)/
targetValue | abs(targetValue) | result | error (A) | abs(Error) |  | result | error (A) | abs(Error) |  | error (A) | error (A) |  | error | specified error
-0.1 | 0.1 | -0.1 | -7.41e-08 | 7.41e-08 |  | -0.099999 | 1.0591e-06 | 1.0591e-06 |  | 2.35e-05 | -2.35e-05 |  | 5.6662e-07 | 2.411%
-0.05 | 0.05 | -0.05 | 1.2842e-07 | 1.2842e-07 |  | -0.05 | 3.3372e-07 | 3.3372e-07 |  | 1.34e-05 | -1.34e-05 |  | 2.3107e-07 | 1.724%
-0.02 | 0.02 | -0.02 | -5.939e-08 | 5.939e-08 |  | -0.02 | 2.9449e-07 | 2.9449e-07 |  | 7.34e-06 | -7.34e-06 |  | 1.7694e-07 | 2.411%
-0.01 | 0.01 | -0.01 | -9.37e-09 | 9.37e-09 |  | -0.01 | 2.8875e-08 | 2.8875e-08 |  | 9.6e-07 | -9.6e-07 |  | 1.9122e-08 | 1.992%
-0.005 | 0.005 | -0.005 | -1.9604e-08 | 1.9604e-08 |  | -0.005 | 1.3694e-08 | 1.3694e-08 |  | 5.9e-07 | -5.9e-07 |  | 1.6649e-08 | 2.822%
-0.002 | 0.002 | -0.002 | -1.8718e-08 | 1.8718e-08 |  | -0.002 | 3.2834e-08 | 3.2834e-08 |  | 3.68e-07 | -3.68e-07 |  | 2.5776e-08 | 7.004%
-0.001 | 0.001 | -0.001 | 2.096e-10 | 2.096e-10 |  | -0.001 | 2.674e-09 | 2.674e-09 |  | 7.2e-08 | -7.2e-08 |  | 1.4418e-09 | 2.002%
-0.0005 | 0.0005 | -0.0005 | 1.3519e-09 | 1.3519e-09 |  | -0.00049999 | 5.0708e-09 | 5.0708e-09 |  | 4.6e-08 | -4.6e-08 |  | 3.2114e-09 | 6.981%
-0.0002 | 0.0002 | -0.0002 | -1.2064e-09 | 1.2064e-09 |  | -0.0002 | 2.9129e-09 | 2.9129e-09 |  | 3.04e-08 | -3.04e-08 |  | 2.0597e-09 | 6.775%
-0.0001 | 0.0001 | -0.0001 | -4.64e-11 | 4.64e-11 |  | -0.0001 | 3.1669e-10 | 3.1669e-10 |  | 8.5e-09 | -8.5e-09 |  | 1.8155e-10 | 2.136%
-5e-05 | 5e-05 | -5e-05 | 6.822e-11 | 6.822e-11 |  | -5e-05 | 4.5832e-10 | 4.5832e-10 |  | 5.9e-09 | -5.9e-09 |  | 2.6327e-10 | 4.462%
-2e-05 | 2e-05 | -2e-05 | -7.891e-11 | 7.891e-11 |  | -2e-05 | 4.5899e-10 | 4.5899e-10 |  | 4.34e-09 | -4.34e-09 |  | 2.6895e-10 | 6.197%
-1e-05 | 1e-05 | -1e-05 | -5.382e-11 | 5.382e-11 |  | -9.9997e-06 | 2.9691e-10 | 2.9691e-10 |  | 3.82e-09 | -3.82e-09 |  | 1.7537e-10 | 4.591%
-5e-06 | 5e-06 | -5e-06 | 2.2461e-11 | 2.2461e-11 |  | -4.9997e-06 | 2.8276e-10 | 2.8276e-10 |  | 3.56e-09 | -3.56e-09 |  | 1.5261e-10 | 4.287%
-2e-06 | 2e-06 | -2e-06 | 3.7955e-11 | 3.7955e-11 |  | -1.9997e-06 | 3.3422e-10 | 3.3422e-10 |  | 3.404e-09 | -3.404e-09 |  | 1.8609e-10 | 5.467%
-1e-06 | 1e-06 | -1.0002e-06 | -2.0799e-10 | 2.0799e-10 |  | -9.9955e-07 | 4.4946e-10 | 4.4946e-10 |  | 3.352e-09 | -3.352e-09 |  | 3.2873e-10 | 9.807%
-5e-07 | 5e-07 | -5.0006e-07 | -6.206e-11 | 6.206e-11 |  | -4.995e-07 | 5.0462e-10 | 5.0462e-10 |  | 3.326e-09 | -3.326e-09 |  | 2.8334e-10 | 8.519%
-2e-07 | 2e-07 | -2.0023e-07 | -2.316e-10 | 2.316e-10 |  | -1.998e-07 | 1.9894e-10 | 1.9894e-10 |  | 3.3104e-09 | -3.3104e-09 |  | 2.1527e-10 | 6.503%
-1e-07 | 1e-07 | -1.0009e-07 | -8.8387e-11 | 8.8387e-11 |  | -9.9739e-08 | 2.6089e-10 | 2.6089e-10 |  | 3.3052e-09 | -3.3052e-09 |  | 1.7464e-10 | 5.284%
-5e-08 | 5e-08 | -5.0006e-08 | -6.2392e-12 | 6.2392e-12 |  | -4.9629e-08 | 3.7113e-10 | 3.7113e-10 |  | 3.3026e-09 | -3.3026e-09 |  | 1.8868e-10 | 5.713%
-2e-08 | 2e-08 | -2.0185e-08 | -1.8472e-10 | 1.8472e-10 |  | -1.9753e-08 | 2.4702e-10 | 2.4702e-10 |  | 3.301e-09 | -3.301e-09 |  | 2.1587e-10 | 6.539%
-1e-08 | 1e-08 | -1.0194e-08 | -1.9404e-10 | 1.9404e-10 |  | -9.7727e-09 | 2.2732e-10 | 2.2732e-10 |  | 3.3005e-09 | -3.3005e-09 |  | 2.1068e-10 | 6.383%
-5e-09 | 5e-09 | -5.171e-09 | -1.7102e-10 | 1.7102e-10 |  | -4.8033e-09 | 1.9668e-10 | 1.9668e-10 |  | 3.3003e-09 | -3.3003e-09 |  | 1.8385e-10 | 5.571%
-2e-09 | 2e-09 | -2.1633e-09 | -1.6333e-10 | 1.6333e-10 |  | -1.8012e-09 | 1.9881e-10 | 1.9881e-10 |  | 3.3001e-09 | -3.3001e-09 |  | 1.8107e-10 | 5.487%
-1e-09 | 1e-09 | -1.1855e-09 | -1.8546e-10 | 1.8546e-10 |  | -7.8507e-10 | 2.1493e-10 | 2.1493e-10 |  | 3.3001e-09 | -3.3001e-09 |  | 2.002e-10 | 6.066%
-5e-10 | 5e-10 | -6.6938e-10 | -1.6938e-10 | 1.6938e-10 |  | -2.8771e-10 | 2.1229e-10 | 2.1229e-10 |  | 3.3e-09 | -3.3e-09 |  | 1.9083e-10 | 5.783%
-2e-10 | 2e-10 | -3.7335e-10 | -1.7335e-10 | 1.7335e-10 |  | -1.2572e-12 | 1.9874e-10 | 1.9874e-10 |  | 3.3e-09 | -3.3e-09 |  | 1.8605e-10 | 5.638%
-1e-10 | 1e-10 | -2.9393e-10 | -1.9393e-10 | 1.9393e-10 |  | 1.06e-10 | 2.06e-10 | 2.06e-10 |  | 3.3e-09 | -3.3e-09 |  | 1.9996e-10 | 6.059%
0 | 0 | -1.8866e-10 | -1.8866e-10 | 1.8866e-10 |  | 2.2059e-10 | 2.2059e-10 | 2.2059e-10 |  | 3.3e-09 | -3.3e-09 |  | 2.0463e-10 | 6.201%
1e-10 | 1e-10 | -7.8807e-11 | -1.7881e-10 | 1.7881e-10 |  | 2.8707e-10 | 1.8707e-10 | 1.8707e-10 |  | 3.3e-09 | -3.3e-09 |  | 1.8294e-10 | 5.544%
2e-10 | 2e-10 | 1.2397e-11 | -1.876e-10 | 1.876e-10 |  | 4.1801e-10 | 2.1801e-10 | 2.1801e-10 |  | 3.3e-09 | -3.3e-09 |  | 2.0281e-10 | 6.146%
5e-10 | 5e-10 | 3.1851e-10 | -1.8149e-10 | 1.8149e-10 |  | 6.6646e-10 | 1.6646e-10 | 1.6646e-10 |  | 3.3e-09 | -3.3e-09 |  | 1.7398e-10 | 5.272%
1e-09 | 1e-09 | 8.2467e-10 | -1.7533e-10 | 1.7533e-10 |  | 1.1955e-09 | 1.9552e-10 | 1.9552e-10 |  | 3.3001e-09 | -3.3001e-09 |  | 1.8543e-10 | 5.619%
2e-09 | 2e-09 | 1.7974e-09 | -2.0257e-10 | 2.0257e-10 |  | 2.2073e-09 | 2.0734e-10 | 2.0734e-10 |  | 3.3001e-09 | -3.3001e-09 |  | 2.0496e-10 | 6.211%
5e-09 | 5e-09 | 4.8061e-09 | -1.939e-10 | 1.939e-10 |  | 5.2021e-09 | 2.0213e-10 | 2.0213e-10 |  | 3.3003e-09 | -3.3003e-09 |  | 1.9802e-10 | 6%
1e-08 | 1e-08 | 9.8461e-09 | -1.5393e-10 | 1.5393e-10 |  | 1.0529e-08 | 5.2929e-10 | 5.2929e-10 |  | 3.3005e-09 | -3.3005e-09 |  | 3.4161e-10 | 10.35%
2e-08 | 2e-08 | 1.9842e-08 | -1.5789e-10 | 1.5789e-10 |  | 2.048e-08 | 4.801e-10 | 4.801e-10 |  | 3.301e-09 | -3.301e-09 |  | 3.19e-10 | 9.664%
5e-08 | 5e-08 | 5.0009e-08 | 8.7791e-12 | 8.7791e-12 |  | 5.0375e-08 | 3.7527e-10 | 3.7527e-10 |  | 3.3026e-09 | -3.3026e-09 |  | 1.9202e-10 | 5.814%
1e-07 | 1e-07 | 9.9953e-08 | -4.6987e-11 | 4.6987e-11 |  | 1.0026e-07 | 2.593e-10 | 2.593e-10 |  | 3.3052e-09 | -3.3052e-09 |  | 1.5314e-10 | 4.633%
2e-07 | 2e-07 | 1.9981e-07 | -1.9281e-10 | 1.9281e-10 |  | 2.0027e-07 | 2.6907e-10 | 2.6907e-10 |  | 3.3104e-09 | -3.3104e-09 |  | 2.3094e-10 | 6.976%
5e-07 | 5e-07 | 4.9994e-07 | -5.7659e-11 | 5.7659e-11 |  | 5.0033e-07 | 3.2513e-10 | 3.2513e-10 |  | 3.326e-09 | -3.326e-09 |  | 1.914e-10 | 5.755%
1e-06 | 1e-06 | 1e-06 | 3.1426e-11 | 3.1426e-11 |  | 1.0004e-06 | 3.959e-10 | 3.959e-10 |  | 3.352e-09 | -3.352e-09 |  | 2.1366e-10 | 6.374%
2e-06 | 2e-06 | 2e-06 | 2.1608e-11 | 2.1608e-11 |  | 2.0005e-06 | 5.0645e-10 | 5.0645e-10 |  | 3.404e-09 | -3.404e-09 |  | 2.6403e-10 | 7.756%
5e-06 | 5e-06 | 5e-06 | 1.7172e-11 | 1.7172e-11 |  | 5.0002e-06 | 2.1839e-10 | 2.1839e-10 |  | 3.56e-09 | -3.56e-09 |  | 1.1778e-10 | 3.308%
1e-05 | 1e-05 | 1e-05 | 6.247e-11 | 6.247e-11 |  | 1.0001e-05 | 5.7097e-10 | 5.7097e-10 |  | 3.82e-09 | -3.82e-09 |  | 3.1672e-10 | 8.291%
2e-05 | 2e-05 | 2e-05 | -2.936e-11 | 2.936e-11 |  | 2.0001e-05 | 5.3003e-10 | 5.3003e-10 |  | 4.34e-09 | -4.34e-09 |  | 2.7969e-10 | 6.445%
5e-05 | 5e-05 | 5e-05 | 1.4244e-10 | 1.4244e-10 |  | 5.0001e-05 | 5.4794e-10 | 5.4794e-10 |  | 5.9e-09 | -5.9e-09 |  | 3.4519e-10 | 5.851%
0.0001 | 0.0001 | 0.0001 | 1.198e-10 | 1.198e-10 |  | 0.0001 | 3.135e-10 | 3.135e-10 |  | 8.5e-09 | -8.5e-09 |  | 2.1665e-10 | 2.549%
0.0002 | 0.0002 | 0.0002 | -8.875e-10 | 8.875e-10 |  | 0.0002 | 4.8003e-09 | 4.8003e-09 |  | 3.04e-08 | -3.04e-08 |  | 2.8439e-09 | 9.355%
0.0005 | 0.0005 | 0.0005 | -7.21e-11 | 7.21e-11 |  | 0.0005 | 3.5442e-09 | 3.5442e-09 |  | 4.6e-08 | -4.6e-08 |  | 1.8082e-09 | 3.931%
0.001 | 0.001 | 0.001 | 2.83e-10 | 2.83e-10 |  | 0.001 | 1.66e-09 | 1.66e-09 |  | 7.2e-08 | -7.2e-08 |  | 9.715e-10 | 1.349%
0.002 | 0.002 | 0.002 | -3.7166e-08 | 3.7166e-08 |  | 0.002 | 2.853e-08 | 2.853e-08 |  | 3.68e-07 | -3.68e-07 |  | 3.2848e-08 | 8.926%
0.005 | 0.005 | 0.005 | -2.853e-08 | 2.853e-08 |  | 0.005 | 4.5151e-08 | 4.5151e-08 |  | 5.9e-07 | -5.9e-07 |  | 3.684e-08 | 6.244%
0.01 | 0.01 | 0.01 | -2.237e-09 | 2.237e-09 |  | 0.01 | 3.035e-08 | 3.035e-08 |  | 9.6e-07 | -9.6e-07 |  | 1.6293e-08 | 1.697%
0.02 | 0.02 | 0.02 | 1.9176e-07 | 1.9176e-07 |  | 0.02 | 1.0126e-07 | 1.0126e-07 |  | 7.34e-06 | -7.34e-06 |  | 1.4651e-07 | 1.996%
0.05 | 0.05 | 0.05 | 4.4292e-07 | 4.4292e-07 |  | 0.05 | 4.7939e-07 | 4.7939e-07 |  | 1.34e-05 | -1.34e-05 |  | 4.6115e-07 | 3.441%
0.1 | 0.1 | 0.1 | 3.996e-07 | 3.996e-07 |  | 0.1 | -2.771e-07 | 2.771e-07 |  | 2.35e-05 | -2.35e-05 |  | 3.3835e-07 | 1.44%
[13 empty] | overall | 5.652%
average