## 3458A drift test (SN18) - 2022-
CAL 72 Temperature Coefficient Adjustment
May/June '22 - 3458A 2823A-01045 - SN18 drift observations | [12 empty] | (From Temperature Regressions)
[6 empty] | Simple Calculations | [5 empty] | Δ Temperature | 0.49324
|  | Raw Data |  |  |  | Delta Time | [5 empty] | Δ Temperature ^3 | -0.001118
[5 empty] | Delta Time | Between |  |  | Temperature Regression Set-up |  |  |  |  | Adjusted
Date/Time | Voltage (V) | Voltage Noise (PPM) | Internal Temperature | CAL 72 Meter | Since Start | Readings | SN18 Drift | Delta Time | Delta Temperature | Delta Temperatue ^3 | Delta CAL 72 PPM |  | Temperature Effect | CAL 72 PPM | Daily Adjusted Drift
2022-05-19 | 10.0001 | 0.026656 | 38.7888 | 0.98694 | 0 |  |  | 0 | 0 | 0 | 0 |  | 0 | 0
2022-05-19 01:01 | 10.0001 | 0.023842 | 38.7313 | 0.98694 | 0.041772 | 1.0025 | -1.7949 | 0.041772 | -0.05751 | -0.00019021 | -0.074979 |  | -0.028366 | -0.046613
2022-05-19 02:01 | 10.0001 | 0.023842 | 38.6569 | 0.98694 | 0.083544 | 1.0025 | -1.5767 | 0.083544 | -0.13189 | -0.0022943 | -0.13172 |  | -0.065051 | -0.066669
2022-05-19 03:01 | 10.0001 | NaN | 38.6055 | 0.98694 | 0.12531 | 1.0024 | -0.89751 | 0.12531 | -0.18331 | -0.00616 | -0.11247 |  | -0.09041 | -0.022058
2022-05-19 04:01 | 10.0001 | 0.064196 | 38.5958 | 0.98694 | 0.16708 | 1.0025 | -1.2371 | 0.16708 | -0.19307 | -0.0071964 | -0.2067 |  | -0.095219 | -0.11148
2022-05-19 05:00 | 10.0001 | 0.026656 | 38.5904 | 0.98694 | 0.20816 | 0.98585 | -0.89076 | 0.20816 | -0.19841 | -0.0078102 | -0.18542 |  | -0.097853 | -0.087568
2022-05-19 06:00 | 10.0001 | 0.023842 | 38.5268 | 0.98694 | 0.24993 | 1.0025 | -0.88379 | 0.24993 | -0.26206 | -0.017996 | -0.22088 |  | -0.12924 | -0.091648
2022-05-19 07:00 | 10.0001 | NaN | 38.4591 | 0.98694 | 0.2917 | 1.0025 | -0.79197 | 0.2917 | -0.32977 | -0.035861 | -0.23102 |  | -0.16261 | -0.068403
2022-05-19 08:01 | 10.0001 | 0.0292 | 38.1541 | 0.98694 | 0.33347 | 1.0025 | -1.2306 | 0.33347 | -0.63476 | -0.25576 | -0.41036 |  | -0.3128 | -0.097557
2022-05-19 09:01 | 10.0001 | 0.023842 | 37.8652 | 0.98694 | 0.37524 | 1.0026 | -1.7119 | 0.37524 | -0.92361 | -0.78788 | -0.64239 |  | -0.45468 | -0.18771
2022-05-19 10:01 | 10.0001 | 0.016859 | 37.3779 | 0.98694 | 0.41702 | 1.0025 | -2.3058 | 0.41702 | -1.4109 | -2.8088 | -0.96155 |  | -0.69279 | -0.26877
2022-05-19 11:00 | 10.0001 | 0.020648 | 37.4688 | 0.98694 | 0.45809 | 0.98587 | -2.1167 | 0.45809 | -1.32 | -2.3001 | -0.96966 |  | -0.64852 | -0.32114
2022-05-19 12:00 | 10.0001 | 0.011921 | 37.9924 | 0.98694 | 0.49986 | 1.0025 | -1.3926 | 0.49986 | -0.79638 | -0.50508 | -0.69609 |  | -0.39224 | -0.30385
2022-05-19 13:00 | 10.0001 | 0.020648 | 38.206 | 0.98694 | 0.54163 | 1.0025 | -1.1355 | 0.54163 | -0.58282 | -0.19797 | -0.61503 |  | -0.28725 | -0.32778
2022-05-19 14:00 | 10.0001 | 0.023842 | 38.3887 | 0.98694 | 0.58341 | 1.0025 | -0.83538 | 0.58341 | -0.40011 | -0.064054 | -0.48736 |  | -0.19728 | -0.29008
2022-05-19 15:01 | 10.0001 | 0.020648 | 38.4683 | 0.98694 | 0.62518 | 1.0025 | -0.80225 | 0.62518 | -0.32048 | -0.032915 | -0.50155 |  | -0.15803 | -0.34351
2022-05-19 16:01 | 10.0001 | 0.0292 | 38.3057 | 0.98694 | 0.66695 | 1.0026 | -0.86442 | 0.66695 | -0.48313 | -0.11277 | -0.57653 |  | -0.23817 | -0.33835
2022-05-19 17:01 | 10.0001 | 0.016859 | 38.6527 | 0.98694 | 0.70872 | 1.0025 | -0.70768 | 0.70872 | -0.13612 | -0.0025218 | -0.50155 |  | -0.067134 | -0.43441
2022-05-19 18:00 | 10.0001 | 0.026656 | 39.0892 | 0.98694 | 0.7498 | 0.98583 | -0.29864 | 0.7498 | 0.30039 | 0.027105 | -0.22392 |  | 0.14813 | -0.37206
2022-05-19 19:00 | 10.0001 | 0.0292 | 39.3197 | 0.98694 | 0.79157 | 1.0025 | -0.25088 | 0.79157 | 0.53091 | 0.14965 | -0.19859 |  | 0.2617 | -0.46029
2022-05-19 20:00 | 10.0001 | 0.020648 | 39.0384 | 0.98694 | 0.83335 | 1.0026 | -0.40974 | 0.83335 | 0.24957 | 0.015545 | -0.34146 |  | 0.12308 | -0.46454
2022-05-19 21:01 | 10.0001 | 0 | 38.8245 | 0.98694 | 0.87512 | 1.0026 | -0.4666 | 0.87512 | 0.035706 | 4.5522e-05 | -0.40833 |  | 0.017611 | -0.42594
2022-05-19 22:01 | 10.0001 | 0.023842 | 38.3453 | 0.98694 | 0.91689 | 1.0024 | -0.74924 | 0.91689 | -0.4435 | -0.087233 | -0.68697 |  | -0.21865 | -0.46832
2022-05-19 23:01 | 10.0001 | 0.0292 | 37.8673 | 0.98694 | 0.95866 | 1.0026 | -0.98505 | 0.95866 | -0.92151 | -0.78254 | -0.94433 |  | -0.45365 | -0.49068
2022-05-20 00:01 | 10.0001 | 0.039537 | 38.0723 | 0.98694 | 1.0004 | 1.0024 | -0.83049 | 1.0004 | -0.71652 | -0.36786 | -0.83085 |  | -0.353 | -0.47785 | -0.47764
2022-05-20 01:00 | 10.0001 | 0.0292 | 38.0034 | 0.98694 | 1.0415 | 0.98578 | -0.87751 | 1.0415 | -0.78542 | -0.48452 | -0.91393 |  | -0.38686 | -0.52707 | -0.48059
2022-05-20 02:00 | 10.0001 | 0.0292 | 37.905 | 0.98694 | 1.0833 | 1.0024 | -0.86893 | 1.0833 | -0.88383 | -0.69041 | -0.94129 |  | -0.43517 | -0.50612 | -0.43957
2022-05-20 03:00 | 10.0001 | 0 | 37.7104 | 0.98694 | 1.125 | 1.0024 | -0.92223 | 1.125 | -1.0784 | -1.2541 | -1.0375 |  | -0.53051 | -0.50704 | -0.48511
2022-05-20 04:01 | 10.0001 | 0.037697 | 37.6124 | 0.98694 | 1.1668 | 1.0024 | -0.91614 | 1.1668 | -1.1764 | -1.6282 | -1.069 |  | -0.57845 | -0.49051 | -0.37914
2022-05-20 05:01 | 10.0001 | 0.011921 | 37.5067 | 0.98694 | 1.2086 | 1.0024 | -0.9373 | 1.2086 | -1.2821 | -2.1076 | -1.1328 |  | -0.63004 | -0.50275 | -0.41501
2022-05-20 06:01 | 10.0001 | 0.050576 | 37.3579 | 0.98694 | 1.2503 | 1.0024 | -0.97162 | 1.2503 | -1.4309 | -2.9297 | -1.2149 |  | -0.7025 | -0.51236 | -0.42054
2022-05-20 07:01 | 10.0001 | 0.033717 | 37.3258 | 0.98694 | 1.2921 | 1.0025 | -0.94962 | 1.2921 | -1.463 | -3.1316 | -1.227 |  | -0.71812 | -0.5089 | -0.44031
2022-05-20 08:00 | 10.0001 | 0.020648 | 37.3145 | 0.98694 | 1.3332 | 0.98572 | -0.95305 | 1.3332 | -1.4744 | -3.2049 | -1.2706 |  | -0.72363 | -0.54696 | -0.44953
2022-05-20 09:00 | 10.0001 | 0.016859 | 36.7536 | 0.98694 | 1.375 | 1.0025 | -1.126 | 1.375 | -2.0352 | -8.4297 | -1.5482 |  | -0.99441 | -0.55381 | -0.36621
2022-05-20 10:00 | 10.0001 | 0.011921 | 36.7073 | 0.98694 | 1.4167 | 1.0025 | -1.1214 | 1.4167 | -2.0816 | -9.0192 | -1.5887 |  | -1.0166 | -0.57212 | -0.30344
2022-05-20 11:01 | 10.0001 | 0.020648 | 36.5423 | 0.98694 | 1.4585 | 1.0026 | -1.1324 | 1.4585 | -2.2465 | -11.3374 | -1.6516 |  | -1.0954 | -0.55619 | -0.23495
2022-05-20 12:01 | 10.0001 | 0 | 36.3676 | 0.98694 | 1.5003 | 1.0026 | -1.1414 | 1.5003 | -2.4212 | -14.1942 | -1.7124 |  | -1.1784 | -0.53398 | -0.23004
2022-05-20 13:01 | 10.0001 | NaN | 36.3304 | 0.98694 | 1.542 | 1.0026 | -1.1308 | 1.542 | -2.4585 | -14.8592 | -1.7438 |  | -1.196 | -0.54777 | -0.2199
2022-05-20 14:00 | 10.0001 | 0.023842 | 36.4417 | 0.98694 | 1.5831 | 0.98595 | -1.1181 | 1.5831 | -2.3471 | -12.9295 | -1.7701 |  | -1.1432 | -0.6269 | -0.33691
2022-05-20 15:00 | 10.0001 | 0.020648 | 36.6497 | 0.98694 | 1.6249 | 1.0026 | -1.0183 | 1.6249 | -2.1391 | -9.7885 | -1.6546 |  | -1.0442 | -0.61044 | -0.267
2022-05-20 16:00 | 10.0001 | 0.011921 | 36.7226 | 0.98694 | 1.6667 | 1.0025 | -0.9733 | 1.6667 | -2.0663 | -8.8217 | -1.6222 |  | -1.0093 | -0.61288 | -0.27461
2022-05-20 17:01 | 10.0001 | 0.011921 | 36.8785 | 0.98694 | 1.7085 | 1.0026 | -0.9412 | 1.7085 | -1.9103 | -6.9717 | -1.608 |  | -0.93446 | -0.67353 | -0.23918
2022-05-20 18:01 | 10.0001 | 0.035763 | 37.0738 | 0.98694 | 1.7502 | 1.0025 | -0.86084 | 1.7502 | -1.715 | -5.0445 | -1.5067 |  | -0.84028 | -0.66639 | -0.29421
2022-05-20 19:01 | 10.0001 | 0 | 37.0357 | 0.98694 | 1.792 | 1.0026 | -0.87018 | 1.792 | -1.7531 | -5.3878 | -1.5594 |  | -0.85867 | -0.70069 | -0.2403
2022-05-20 20:01 | 10.0001 | 0.016859 | 36.8806 | 0.98694 | 1.8338 | 1.0025 | -0.8592 | 1.8338 | -1.9082 | -6.9487 | -1.5756 |  | -0.93345 | -0.64212 | -0.17751
2022-05-20 21:00 | 10.0001 | 0.033717 | 36.7015 | 0.98694 | 1.8749 | 0.98594 | -0.91765 | 1.8749 | -2.0874 | -9.0947 | -1.7205 |  | -1.0194 | -0.70107 | -0.2752
2022-05-20 22:00 | 10.0001 | 0.023842 | 36.5759 | 0.98694 | 1.9166 | 1.0025 | -0.91933 | 1.9166 | -2.2129 | -10.8368 | -1.762 |  | -1.0794 | -0.68262 | -0.21436
2022-05-20 23:00 | 10.0001 | 0.023842 | 36.4779 | 0.98694 | 1.9584 | 1.0026 | -0.94525 | 1.9584 | -2.311 | -12.342 | -1.8512 |  | -1.1261 | -0.72511 | -0.23449
2022-05-21 00:01 | 10.0001 | 0.016859 | 36.2678 | 0.98694 | 2.0002 | 1.0025 | -0.95894 | 2.0002 | -2.521 | -16.022 | -1.918 |  | -1.2255 | -0.6925 | -0.21471
2022-05-21 01:01 | 10.0001 | 0.026656 | 36.1486 | 0.98694 | 2.0419 | 1.0026 | -0.97704 | 2.0419 | -2.6402 | -18.4039 | -1.995 |  | -1.2817 | -0.71338 | -0.18622
2022-05-21 02:01 | 10.0001 | 0.023842 | 36.0655 | 0.98694 | 2.0837 | 1.0026 | -0.97787 | 2.0837 | -2.7234 | -20.1986 | -2.0376 |  | -1.3207 | -0.71691 | -0.2107
2022-05-21 03:00 | 10.0001 | 0.020648 | 35.9929 | 0.98694 | 2.1248 | 0.9859 | -1.0081 | 2.1248 | -2.796 | -21.8573 | -2.142 |  | -1.3546 | -0.78732 | -0.28035
2022-05-21 04:00 | 10.0001 | 0.023842 | 35.8014 | 0.98694 | 2.1666 | 1.0026 | -0.99472 | 2.1666 | -2.9874 | -26.6606 | -2.1551 |  | -1.4437 | -0.71146 | -0.221
2022-05-21 05:00 | 10.0001 | 0.026656 | 35.7024 | 0.98694 | 2.2083 | 1.0025 | -1.0057 | 2.2083 | -3.0864 | -29.4002 | -2.221 |  | -1.4895 | -0.73154 | -0.22885
2022-05-21 06:01 | 10.0001 | 0.011921 | 35.6176 | 0.98694 | 2.2501 | 1.0026 | -1.0208 | 2.2501 | -3.1712 | -31.8915 | -2.297 |  | -1.5285 | -0.76848 | -0.25618
2022-05-21 07:01 | 10.0001 | 0.026656 | 35.5871 | 0.98694 | 2.2919 | 1.0026 | -0.99206 | 2.2919 | -3.2018 | -32.8223 | -2.2737 |  | -1.5425 | -0.73115 | -0.2223
2022-05-21 08:01 | 10.0001 | 0.023842 | 35.6307 | 0.98694 | 2.3337 | 1.0026 | -0.95346 | 2.3337 | -3.1581 | -31.499 | -2.2251 |  | -1.5225 | -0.70255 | -0.15551
2022-05-21 09:01 | 10.0001 | 0.023842 | 35.9061 | 0.98694 | 2.3754 | 1.0026 | -0.93115 | 2.3754 | -2.8827 | -23.9562 | -2.2119 |  | -1.3951 | -0.81678 | -0.26285
2022-05-21 10:00 | 10.0001 | 0.026656 | 36.2295 | 0.98694 | 2.4165 | 0.98591 | -0.8411 | 2.4165 | -2.5593 | -16.7633 | -2.0325 |  | -1.2436 | -0.78894 | -0.21686
2022-05-21 11:00 | 10.0001 | 0.026656 | 36.6109 | 0.98694 | 2.4583 | 1.0025 | -0.75468 | 2.4583 | -2.1779 | -10.3301 | -1.8552 |  | -1.0627 | -0.79256 | -0.23642
2022-05-21 12:00 | 10.0001 | 0 | 36.8469 | 0.98694 | 2.5001 | 1.0026 | -0.70681 | 2.5001 | -1.942 | -7.3236 | -1.7671 |  | -0.94966 | -0.81741 | -0.28349
2022-05-21 13:01 | 10.0001 | 0 | 37.1826 | 0.98694 | 2.5418 | 1.0026 | -0.64497 | 2.5418 | -1.6063 | -4.1443 | -1.6394 |  | -0.78764 | -0.85176 | -0.30406
2022-05-21 14:01 | 10.0001 | 0.033717 | 37.2907 | 0.98694 | 2.5836 | 1.0026 | -0.61964 | 2.5836 | -1.4981 | -3.3622 | -1.6009 |  | -0.73516 | -0.86574 | -0.23873
2022-05-21 15:01 | 10.0001 | 0.016859 | 37.3867 | 0.98694 | 2.6254 | 1.0026 | -0.58469 | 2.6254 | -1.4022 | -2.7568 | -1.535 |  | -0.68852 | -0.84652 | -0.23597
2022-05-21 16:00 | 10.0001 | 0.035763 | 37.4512 | 0.98694 | 2.6665 | 0.98583 | -0.56542 | 2.6665 | -1.3376 | -2.3933 | -1.5077 |  | -0.65709 | -0.8506 | -0.23776
2022-05-21 17:00 | 10.0001 | 0.03154 | 37.5393 | 0.98694 | 2.7082 | 1.0025 | -0.5567 | 2.7082 | -1.2495 | -1.9508 | -1.5077 |  | -0.61412 | -0.89357 | -0.22008
2022-05-21 18:00 | 10.0001 | 0.011921 | 37.5453 | 0.98694 | 2.75 | 1.0025 | -0.56556 | 2.75 | -1.2436 | -1.9231 | -1.5553 |  | -0.61122 | -0.94409 | -0.27776
2022-05-21 19:01 | 10.0001 | 0.016859 | 37.6494 | 0.98694 | 2.7918 | 1.0026 | -0.53641 | 2.7918 | -1.1394 | -1.4794 | -1.4976 |  | -0.56036 | -0.93719 | -0.23655
2022-05-21 20:01 | 10.0001 | 0.011921 | 37.7347 | 0.98694 | 2.8336 | 1.0026 | -0.49954 | 2.8336 | -1.0541 | -1.1712 | -1.4155 |  | -0.51861 | -0.89687 | -0.2548
2022-05-21 21:01 | 10.0001 | 0.026656 | 37.7109 | 0.98694 | 2.8753 | 1.0025 | -0.5018 | 2.8753 | -1.078 | -1.2526 | -1.4428 |  | -0.53029 | -0.91254 | -0.21137
2022-05-21 22:01 | 10.0001 | 0.03154 | 37.5547 | 0.98694 | 2.9171 | 1.0025 | -0.52275 | 2.9171 | -1.2341 | -1.8795 | -1.5249 |  | -0.60661 | -0.9183 | -0.23557
2022-05-21 23:00 | 10.0001 | 0.035763 | 37.2152 | 0.98694 | 2.9582 | 0.98587 | -0.57372 | 2.9582 | -1.5736 | -3.8967 | -1.6972 |  | -0.77181 | -0.92535 | -0.20028
2022-05-22 | 10.0001 | 0.011921 | 37.0558 | 0.98694 | 3 | 1.0026 | -0.58126 | 3 | -1.7331 | -5.2052 | -1.7438 |  | -0.84899 | -0.89478 | -0.20232
2022-05-22 01:00 | 10.0001 | 0 | 36.7976 | 0.98694 | 3.0417 | 1.0026 | -0.61559 | 3.0417 | -1.9913 | -7.8957 | -1.8724 |  | -0.97335 | -0.8991 | -0.18576
2022-05-22 02:01 | 10.0001 | 0.023842 | 36.6997 | 0.98694 | 3.0835 | 1.0026 | -0.62368 | 3.0835 | -2.0891 | -9.1174 | -1.9231 |  | -1.0202 | -0.90289 | -0.18601
2022-05-22 03:01 | 10.0001 | 0.033717 | 36.6228 | 0.98694 | 3.1253 | 1.0025 | -0.62831 | 3.1253 | -2.166 | -10.1622 | -1.9636 |  | -1.057 | -0.90663 | -0.11925
2022-05-22 04:01 | 10.0001 | 0.011921 | 36.4937 | 0.98694 | 3.1671 | 1.0025 | -0.64274 | 3.1671 | -2.2951 | -12.09 | -2.0356 |  | -1.1185 | -0.91704 | -0.20549
2022-05-22 05:00 | 10.0001 | NaN | 36.4351 | 0.98694 | 3.2081 | 0.98583 | -0.66293 | 3.2081 | -2.3538 | -13.0404 | -2.1268 |  | -1.1464 | -0.98038 | -0.24889
2022-05-22 06:00 | 10.0001 | 0.016859 | 36.3159 | 0.98694 | 3.2499 | 1.0025 | -0.68496 | 3.2499 | -2.473 | -15.1236 | -2.2261 |  | -1.2029 | -1.0232 | -0.25478
2022-05-22 07:00 | 10.0001 | 0.023842 | 36.3593 | 0.98694 | 3.2917 | 1.0025 | -0.66211 | 3.2917 | -2.4295 | -14.3401 | -2.1795 |  | -1.1823 | -0.99717 | -0.26607
2022-05-22 08:01 | 10.0001 | 0.03154 | 36.5324 | 0.98694 | 3.3334 | 1.0026 | -0.63406 | 3.3334 | -2.2564 | -11.4883 | -2.1136 |  | -1.1001 | -1.0135 | -0.31101
2022-05-22 09:01 | 10.0001 | 0.020648 | 36.8528 | 0.98694 | 3.3752 | 1.0026 | -0.58808 | 3.3752 | -1.936 | -7.2566 | -1.9849 |  | -0.94681 | -1.0381 | -0.22137
2022-05-22 10:01 | 10.0001 | 0.020648 | 37.2707 | 0.98694 | 3.417 | 1.0026 | -0.53049 | 3.417 | -1.5182 | -3.499 | -1.8127 |  | -0.7449 | -1.0678 | -0.27869
2022-05-22 11:01 | 10.0001 | 0 | 37.6773 | 0.98694 | 3.4588 | 1.0026 | -0.46227 | 3.4588 | -1.1115 | -1.3733 | -1.5989 |  | -0.54671 | -1.0522 | -0.25948
2022-05-22 12:00 | 10.0001 | 0 | 38.0393 | 0.98694 | 3.4999 | 0.98593 | -0.41139 | 3.4999 | -0.74949 | -0.42102 | -1.4398 |  | -0.36921 | -1.0706 | -0.25323
2022-05-22 13:00 | 10.0001 | 0.020648 | 38.3381 | 0.98694 | 3.5416 | 1.0026 | -0.3928 | 3.5416 | -0.45074 | -0.091573 | -1.3912 |  | -0.22222 | -1.1689 | -0.31725
2022-05-22 14:00 | 10.0001 | NaN | 38.1908 | 0.98694 | 3.5834 | 1.0025 | -0.38624 | 3.5834 | -0.59801 | -0.21385 | -1.3841 |  | -0.29472 | -1.0894 | -0.22366
2022-05-22 15:01 | 10.0001 | 0.016859 | 38.1211 | 0.98694 | 3.6252 | 1.0025 | -0.4008 | 3.6252 | -0.66775 | -0.29774 | -1.453 |  | -0.32903 | -1.1239 | -0.27748
2022-05-22 16:01 | 10.0001 | 0.016859 | 38.1241 | 0.98694 | 3.6669 | 1.0026 | -0.40204 | 3.6669 | -0.66474 | -0.29374 | -1.4742 |  | -0.32755 | -1.1467 | -0.29596
2022-05-22 17:01 | 10.0001 | 0.011921 | 38.3148 | 0.98694 | 3.7087 | 1.0026 | -0.38002 | 3.7087 | -0.47401 | -0.1065 | -1.4094 |  | -0.23368 | -1.1757 | -0.28202
2022-05-22 18:00 | 10.0001 | 0.03154 | 38.522 | 0.98694 | 3.7498 | 0.98585 | -0.35505 | 3.7498 | -0.26683 | -0.018997 | -1.3314 |  | -0.13159 | -1.1998 | -0.25576
2022-05-22 19:00 | 10.0001 | 0.020648 | 38.5188 | 0.98694 | 3.7916 | 1.0026 | -0.34981 | 3.7916 | -0.26998 | -0.019679 | -1.3263 |  | -0.13314 | -1.1932 | -0.25604
2022-05-22 20:00 | 10.0001 | 0.020648 | 38.674 | 0.98694 | 3.8333 | 1.0026 | -0.34282 | 3.8333 | -0.11482 | -0.0015138 | -1.3142 |  | -0.056633 | -1.2575 | -0.36073
2022-05-22 21:01 | 10.0001 | 0.026656 | 38.5933 | 0.98694 | 3.8751 | 1.0026 | -0.33756 | 3.8751 | -0.19552 | -0.0074747 | -1.3081 |  | -0.096431 | -1.2116 | -0.29917
2022-05-22 22:01 | 10.0001 | 0.023842 | 38.4424 | 0.98694 | 3.9169 | 1.0025 | -0.35284 | 3.9169 | -0.3464 | -0.041565 | -1.382 |  | -0.17081 | -1.2112 | -0.29299
2022-05-22 23:01 | 10.0001 | 0.020648 | 38.1261 | 0.98694 | 3.9587 | 1.0026 | -0.39289 | 3.9587 | -0.66273 | -0.29107 | -1.5553 |  | -0.32656 | -1.2288 | -0.30326
2022-05-23 00:01 | 10.0001 | 0.016859 | 37.8516 | 0.98694 | 4.0004 | 1.0026 | -0.42374 | 4.0004 | -0.93724 | -0.82328 | -1.6951 |  | -0.46136 | -1.2338 | -0.33883
2022-05-23 01:00 | 10.0001 | 0.0292 | 37.6866 | 0.98694 | 4.0415 | 0.98591 | -0.44325 | 4.0415 | -1.1022 | -1.3391 | -1.7914 |  | -0.54216 | -1.2492 | -0.3502
2022-05-23 02:00 | 10.0001 | 0.020648 | 37.5595 | 0.98694 | 4.0833 | 1.0025 | -0.46874 | 4.0833 | -1.2293 | -1.8577 | -1.914 |  | -0.60426 | -1.3097 | -0.40693
2022-05-23 03:00 | 10.0001 | 0.011921 | 37.485 | 0.98694 | 4.1251 | 1.0025 | -0.47603 | 4.1251 | -1.3038 | -2.2162 | -1.9636 |  | -0.6406 | -1.323 | -0.4165
2022-05-23 04:01 | 10.0001 | 0.020648 | 37.3183 | 0.98694 | 4.1668 | 1.0026 | -0.48657 | 4.1668 | -1.4706 | -3.1802 | -2.0275 |  | -0.72179 | -1.3057 | -0.38872
2022-05-23 05:01 | 10.0001 | 0.016859 | 37.2133 | 0.98694 | 4.2087 | 1.0037 | -0.4827 | 4.2087 | -1.5755 | -3.9108 | -2.0315 |  | -0.77273 | -1.2588 | -0.27827
2022-05-23 06:01 | 10.0001 | 0 | 37.0928 | 0.98694 | 4.2504 | 1.0026 | -0.50799 | 4.2504 | -1.696 | -4.8783 | -2.1592 |  | -0.83107 | -1.3281 | -0.30475
2022-05-23 07:00 | 10.0001 | 0.026656 | 37.0252 | 0.98694 | 4.2915 | 0.9859 | -0.50408 | 4.2915 | -1.7636 | -5.4855 | -2.1632 |  | -0.86375 | -1.2995 | -0.30237
2022-05-23 08:00 | 10.0001 | 0.011921 | 36.957 | 0.98694 | 4.3333 | 1.0026 | -0.50459 | 4.3333 | -1.8318 | -6.1464 | -2.1865 |  | -0.89663 | -1.2899 | -0.27647
2022-05-23 09:00 | 10.0001 | 0.033717 | 37.1421 | 0.98694 | 4.3751 | 1.0026 | -0.49352 | 4.3751 | -1.6467 | -4.465 | -2.1592 |  | -0.80721 | -1.352 | -0.31392
2022-05-23 10:01 | 10.0001 | 0.016859 | 37.294 | 0.98694 | 4.4168 | 1.0026 | -0.4611 | 4.4168 | -1.4948 | -3.3401 | -2.0366 |  | -0.73357 | -1.303 | -0.2353
2022-05-23 11:01 | 10.0001 | 0.023842 | 37.4877 | 0.98694 | 4.4586 | 1.0025 | -0.44246 | 4.4586 | -1.3012 | -2.2028 | -1.9728 |  | -0.63931 | -1.3334 | -0.28133
2022-05-23 12:01 | 10.0001 | 0.03154 | 37.5645 | 0.98694 | 4.5004 | 1.0025 | -0.4271 | 4.5004 | -1.2243 | -1.8351 | -1.9221 |  | -0.60181 | -1.3203 | -0.24956
2022-05-23 13:00 | 10.0001 | 0.020648 | 37.5922 | 0.98694 | 4.5415 | 0.98594 | -0.42323 | 4.5415 | -1.1966 | -1.7133 | -1.9221 |  | -0.58829 | -1.3338 | -0.16489
2022-05-23 14:00 | 10.0001 | 0.020648 | 37.5695 | 0.98694 | 4.5832 | 1.0026 | -0.43131 | 4.5832 | -1.2193 | -1.8129 | -1.9768 |  | -0.59939 | -1.3774 | -0.28811
2022-05-23 15:00 | 10.0001 | 0.03154 | 37.5449 | 0.98694 | 4.625 | 1.0026 | -0.42457 | 4.625 | -1.2439 | -1.9246 | -1.9636 |  | -0.61138 | -1.3523 | -0.22835
2022-05-23 16:01 | 10.0001 | 0.0292 | 37.613 | 0.98694 | 4.6668 | 1.0025 | -0.41252 | 4.6668 | -1.1758 | -1.6255 | -1.9251 |  | -0.57813 | -1.347 | -0.20034
2022-05-23 17:01 | 10.0001 | 0.020648 | 37.6213 | 0.98694 | 4.7086 | 1.0026 | -0.40456 | 4.7086 | -1.1675 | -1.5915 | -1.9049 |  | -0.57409 | -1.3308 | -0.15508
2022-05-23 18:01 | 10.0001 | 0.026656 | 37.7206 | 0.98694 | 4.7503 | 1.0026 | -0.39353 | 4.7503 | -1.0682 | -1.2189 | -1.8694 |  | -0.52553 | -1.3439 | -0.14401
2022-05-23 19:01 | 10.0001 | 0.026656 | 37.8066 | 0.98694 | 4.7921 | 1.0026 | -0.38143 | 4.7921 | -0.98221 | -0.94756 | -1.8279 |  | -0.4834 | -1.3445 | -0.15121
2022-05-23 20:00 | 10.0001 | 0.023842 | 37.6666 | 0.98694 | 4.8332 | 0.98591 | -0.40733 | 4.8332 | -1.1222 | -1.4132 | -1.9687 |  | -0.55193 | -1.4168 | -0.15927
2022-05-23 21:00 | 10.0001 | 0.023842 | 37.5306 | 0.98694 | 4.875 | 1.0026 | -0.41195 | 4.875 | -1.2582 | -1.9919 | -2.0082 |  | -0.61837 | -1.3898 | -0.17823
2022-05-23 22:00 | 10.0001 | 0.016859 | 37.4017 | 0.98694 | 4.9167 | 1.0025 | -0.42864 | 4.9167 | -1.3872 | -2.6693 | -2.1075 |  | -0.68122 | -1.4263 | -0.2151
2022-05-23 23:01 | 10.0001 | 0 | 37.2663 | 0.98694 | 4.9585 | 1.0025 | -0.43464 | 4.9585 | -1.5225 | -3.5293 | -2.1551 |  | -0.74702 | -1.4081 | -0.1794
2022-05-24 00:01 | 10.0001 | 0.026656 | 37.032 | 0.98694 | 5.0003 | 1.0026 | -0.44681 | 5.0003 | -1.7569 | -5.4227 | -2.2342 |  | -0.86049 | -1.3737 | -0.13993
2022-05-24 01:01 | 10.0001 | 0.020648 | 36.8231 | 0.98694 | 5.042 | 1.0026 | -0.47164 | 5.042 | -1.9657 | -7.5955 | -2.378 |  | -0.96107 | -1.417 | -0.16767
2022-05-24 02:00 | 10.0001 | 0.020648 | 36.6356 | 0.98694 | 5.0831 | 0.98584 | -0.49295 | 5.0831 | -2.1532 | -9.9829 | -2.5057 |  | -1.0509 | -1.4548 | -0.14513
2022-05-24 03:00 | 10.0001 | 0.020648 | 36.5482 | 0.98694 | 5.1249 | 1.0025 | -0.49901 | 5.1249 | -2.2407 | -11.2495 | -2.5574 |  | -1.0926 | -1.4648 | -0.14177
2022-05-24 04:00 | 10.0001 | 0.023842 | 36.441 | 0.98694 | 5.1667 | 1.0025 | -0.50773 | 5.1667 | -2.3478 | -12.9411 | -2.6233 |  | -1.1435 | -1.4797 | -0.17405
2022-05-24 05:01 | 10.0001 | 0 | 36.3155 | 0.98694 | 5.2084 | 1.0026 | -0.50404 | 5.2084 | -2.4733 | -15.1294 | -2.6253 |  | -1.203 | -1.4223 | -0.16352
2022-05-24 06:01 | 10.0001 | 0.03154 | 36.1825 | 0.98694 | 5.2502 | 1.0026 | -0.52454 | 5.2502 | -2.6063 | -17.7041 | -2.754 |  | -1.2657 | -1.4882 | -0.16014
2022-05-24 07:01 | 10.0001 | 0.03154 | 36.104 | 0.98694 | 5.292 | 1.0026 | -0.52346 | 5.292 | -2.6848 | -19.3529 | -2.7702 |  | -1.3026 | -1.4675 | -0.16798
2022-05-24 08:01 | 10.0001 | 0.016859 | 36.0597 | 0.98694 | 5.3338 | 1.0026 | -0.52183 | 5.3338 | -2.7291 | -20.3264 | -2.7833 |  | -1.3234 | -1.46 | -0.16997
2022-05-24 09:00 | 10.0001 | 0.011921 | 36.1651 | 0.98694 | 5.3748 | 0.98594 | -0.50654 | 5.3748 | -2.6237 | -18.0615 | -2.7225 |  | -1.2739 | -1.4486 | -0.096661
2022-05-24 10:00 | 10.0001 | 0.03154 | 36.2939 | 0.98694 | 5.4166 | 1.0025 | -0.50469 | 5.4166 | -2.4949 | -15.5297 | -2.7337 |  | -1.2132 | -1.5205 | -0.21749
2022-05-24 11:00 | 10.0001 | 0.023842 | 36.4364 | 0.98694 | 5.4584 | 1.0025 | -0.49043 | 5.4584 | -2.3524 | -13.0184 | -2.677 |  | -1.1458 | -1.5312 | -0.19779
2022-05-24 12:01 | 10.0001 | 0.020648 | 36.7066 | 0.98694 | 5.5002 | 1.0026 | -0.44949 | 5.5002 | -2.0823 | -9.0284 | -2.4723 |  | -1.017 | -1.4553 | -0.13507
2022-05-24 13:01 | 10.0001 | 0.026656 | 36.7815 | 0.98694 | 5.5419 | 1.0026 | -0.45378 | 5.5419 | -2.0073 | -8.0878 | -2.5148 |  | -0.98103 | -1.5338 | -0.1999
2022-05-24 14:01 | 10.0001 | 0.033717 | 36.7383 | 0.98694 | 5.5837 | 1.0026 | -0.45384 | 5.5837 | -2.0505 | -8.6218 | -2.5341 |  | -1.0018 | -1.5323 | -0.15484
2022-05-24 15:00 | 10.0001 | 0.035763 | 36.8031 | 0.98694 | 5.6248 | 0.98594 | -0.44566 | 5.6248 | -1.9857 | -7.8294 | -2.5067 |  | -0.97066 | -1.5361 | -0.18385
2022-05-24 16:00 | 10.0001 | NaN | 36.7848 | 0.98694 | 5.6666 | 1.0026 | -0.44791 | 5.6666 | -2.004 | -8.048 | -2.5381 |  | -0.97945 | -1.5587 | -0.21173
2022-05-24 17:00 | 10.0001 | 0.023842 | 36.9305 | 0.98694 | 5.7083 | 1.0026 | -0.43505 | 5.7083 | -1.8583 | -6.4171 | -2.4834 |  | -0.9094 | -1.574 | -0.2433
2022-05-24 18:01 | 10.0001 | 0.0292 | 37.0841 | 0.98694 | 5.7501 | 1.0025 | -0.41762 | 5.7501 | -1.7047 | -4.954 | -2.4014 |  | -0.83529 | -1.5661 | -0.22223
2022-05-24 19:01 | 10.0001 | 0.020648 | 37.1343 | 0.98694 | 5.7919 | 1.0025 | -0.4202 | 5.7919 | -1.6545 | -4.5289 | -2.4338 |  | -0.81099 | -1.6228 | -0.27838
2022-05-24 20:01 | 10.0001 | 0.03154 | 37.0334 | 0.98694 | 5.8337 | 1.0025 | -0.42345 | 5.8337 | -1.7554 | -5.4089 | -2.4703 |  | -0.85977 | -1.6105 | -0.19362
2022-05-24 21:01 | 10.0001 | 0.026656 | 36.9775 | 0.98694 | 5.8754 | 1.0025 | -0.42078 | 5.8754 | -1.8113 | -5.9425 | -2.4723 |  | -0.88676 | -1.5855 | -0.19558
2022-05-24 22:00 | 10.0001 | 0.03154 | 36.8098 | 0.98694 | 5.9165 | 0.98589 | -0.43173 | 5.9165 | -1.979 | -7.7506 | -2.5544 |  | -0.96745 | -1.5869 | -0.16064
2022-05-24 23:00 | 10.0001 | 0.020648 | 36.63 | 0.98694 | 5.9583 | 1.0025 | -0.44877 | 5.9583 | -2.1588 | -10.0615 | -2.6739 |  | -1.0536 | -1.6203 | -0.21226
2022-05-25 | 10.0001 | 0.033717 | 36.5335 | 0.98694 | 6.0001 | 1.0026 | -0.45274 | 6.0001 | -2.2553 | -11.4715 | -2.7165 |  | -1.0996 | -1.6169 | -0.24326
2022-05-25 01:01 | 10.0001 | 0.020648 | 36.3137 | 0.98694 | 6.0418 | 1.0025 | -0.46806 | 6.0418 | -2.4751 | -15.1627 | -2.8279 |  | -1.2039 | -1.6241 | -0.20713
2022-05-25 02:01 | 10.0001 | 0.0292 | 36.1351 | 0.98694 | 6.0836 | 1.0026 | -0.47367 | 6.0836 | -2.6537 | -18.6887 | -2.8816 |  | -1.288 | -1.5936 | -0.13869
2022-05-25 03:01 | 10.0001 | NaN | 35.9686 | 0.98694 | 6.1254 | 1.0026 | -0.48632 | 6.1254 | -2.8202 | -22.4314 | -2.9789 |  | -1.366 | -1.6129 | -0.14807
2022-05-25 04:00 | 10.0001 | 0.026656 | 35.8679 | 0.98694 | 6.1665 | 0.98583 | -0.49015 | 6.1665 | -2.9209 | -24.9203 | -3.0225 |  | -1.4128 | -1.6096 | -0.12994
2022-05-25 05:00 | 10.0001 | 0.026656 | 35.7549 | 0.98694 | 6.2083 | 1.0038 | -0.49696 | 6.2083 | -3.0339 | -27.927 | -3.0853 |  | -1.4652 | -1.62 | -0.1978
2022-05-25 06:00 | 10.0001 | 0.03154 | 35.6044 | 0.98694 | 6.2501 | 1.0026 | -0.4985 | 6.2501 | -3.1844 | -32.2926 | -3.1157 |  | -1.5346 | -1.5811 | -0.092883
2022-05-25 07:01 | 10.0001 | 0.011921 | 35.6297 | 0.98694 | 6.2918 | 1.0025 | -0.49777 | 6.2918 | -3.1592 | -31.5292 | -3.1319 |  | -1.523 | -1.6089 | -0.1414
2022-05-25 08:01 | 10.0001 | 0 | 35.7299 | 0.98694 | 6.3336 | 1.0038 | -0.48809 | 6.3336 | -3.0589 | -28.6231 | -3.0914 |  | -1.4768 | -1.6146 | -0.15462
2022-05-25 09:01 | 10.0001 | 0.033717 | 35.7523 | 0.98694 | 6.3754 | 1.0026 | -0.49856 | 6.3754 | -3.0365 | -27.9974 | -3.1785 |  | -1.4664 | -1.7121 | -0.26331
2022-05-25 10:00 | 10.0001 | 0 | 35.8897 | 0.98694 | 6.4165 | 0.98585 | -0.48447 | 6.4165 | -2.8991 | -24.3674 | -3.1086 |  | -1.4027 | -1.7059 | -0.18541
2022-05-25 11:00 | 10.0001 | 0.042982 | 36.0899 | 0.98694 | 6.4583 | 1.0026 | -0.46423 | 6.4583 | -2.6989 | -19.6597 | -2.9981 |  | -1.3092 | -1.6889 | -0.15773
2022-05-25 12:00 | 10.0001 | 0.023842 | 36.0906 | 0.98694 | 6.5 | 1.0026 | -0.45424 | 6.5 | -2.6983 | -19.6451 | -2.9525 |  | -1.3089 | -1.6436 | -0.18833
2022-05-25 13:01 | 10.0001 | 0.023842 | 36.2789 | 0.98694 | 6.5418 | 1.0026 | -0.44235 | 6.5418 | -2.5099 | -15.8119 | -2.8938 |  | -1.2203 | -1.6735 | -0.13968
2022-05-25 14:01 | 10.0001 | 0.03154 | 36.3952 | 0.98694 | 6.5836 | 1.0026 | -0.44478 | 6.5836 | -2.3936 | -13.7133 | -2.9282 |  | -1.1653 | -1.763 | -0.23066
2022-05-25 15:01 | 10.0001 | 0.026656 | 36.5075 | 0.98694 | 6.6254 | 1.0026 | -0.43448 | 6.6254 | -2.2814 | -11.8735 | -2.8786 |  | -1.112 | -1.7666 | -0.2304
2022-05-25 16:00 | 10.0001 | 0 | 36.5975 | 0.98694 | 6.6665 | 0.98592 | -0.42207 | 6.6665 | -2.1913 | -10.5219 | -2.8137 |  | -1.0691 | -1.7447 | -0.18601
2022-05-25 17:00 | 10.0001 | 0.026656 | 36.5733 | 0.98694 | 6.7082 | 1.0026 | -0.41809 | 6.7082 | -2.2155 | -10.875 | -2.8046 |  | -1.0806 | -1.724 | -0.14999
2022-05-25 18:00 | 10.0001 | 0.03154 | 36.7602 | 0.98694 | 6.75 | 1.0025 | -0.4113 | 6.75 | -2.0286 | -8.3484 | -2.7762 |  | -0.99126 | -1.785 | -0.21895
2022-05-25 19:01 | 10.0001 | 0 | 36.9386 | 0.98694 | 6.7918 | 1.0026 | -0.40235 | 6.7918 | -1.8502 | -6.3336 | -2.7327 |  | -0.9055 | -1.8272 | -0.20441
2022-05-25 20:01 | 10.0001 | 0.011921 | 36.8043 | 0.98694 | 6.8335 | 1.0026 | -0.40019 | 6.8335 | -1.9845 | -7.816 | -2.7347 |  | -0.97012 | -1.7646 | -0.15413
2022-05-25 21:01 | 10.0001 | 0.03154 | 36.6216 | 0.98694 | 6.8753 | 1.0026 | -0.40586 | 6.8753 | -2.1673 | -10.1796 | -2.7904 |  | -1.0576 | -1.7328 | -0.14733
2022-05-25 22:01 | 10.0001 | 0.03154 | 36.5108 | 0.98694 | 6.9171 | 1.0025 | -0.41659 | 6.9171 | -2.2781 | -11.822 | -2.8816 |  | -1.1104 | -1.7712 | -0.18421
2022-05-25 23:00 | 10.0001 | 0.0292 | 36.3697 | 0.98694 | 6.9582 | 0.98593 | -0.41705 | 6.9582 | -2.4191 | -14.1566 | -2.9019 |  | -1.1774 | -1.7245 | -0.1042
2022-05-26 | 10.0001 | 0.020648 | 36.3156 | 0.98694 | 6.9999 | 1.0026 | -0.43063 | 6.9999 | -2.4732 | -15.1283 | -3.0144 |  | -1.203 | -1.8114 | -0.19452
2022-05-26 01:00 | 10.0001 | 0.037697 | 36.1171 | 0.98694 | 7.0417 | 1.0026 | -0.44045 | 7.0417 | -2.6717 | -19.0703 | -3.1015 |  | -1.2965 | -1.805 | -0.18099
2022-05-26 02:01 | 10.0001 | 0.0292 | 35.9806 | 0.98694 | 7.0835 | 1.0025 | -0.43885 | 7.0835 | -2.8082 | -22.1462 | -3.1086 |  | -1.3604 | -1.7482 | -0.15465
2022-05-26 03:01 | 10.0001 | 0.020648 | 35.8883 | 0.98694 | 7.1253 | 1.0026 | -0.44524 | 7.1253 | -2.9005 | -24.4012 | -3.1724 |  | -1.4033 | -1.7691 | -0.15616
2022-05-26 04:01 | 10.0001 | NaN | 35.7938 | 0.98694 | 7.167 | 1.0025 | -0.45367 | 7.167 | -2.995 | -26.8651 | -3.2515 |  | -1.4472 | -1.8042 | -0.19451
2022-05-26 05:00 | 10.0001 | 0.016859 | 35.7129 | 0.98694 | 7.2081 | 0.98598 | -0.46078 | 7.2081 | -3.0759 | -29.1019 | -3.3214 |  | -1.4846 | -1.8367 | -0.21673
2022-05-26 06:00 | 10.0001 | 0.051962 | 35.6611 | 0.98694 | 7.2499 | 1.0026 | -0.45533 | 7.2499 | -3.1277 | -30.5969 | -3.3011 |  | -1.5085 | -1.7926 | -0.21155
2022-05-26 07:00 | 10.0001 | NaN | 35.6973 | 0.98694 | 7.2917 | 1.0026 | -0.46287 | 7.2917 | -3.0915 | -29.5476 | -3.3751 |  | -1.4918 | -1.8832 | -0.27436
2022-05-26 08:01 | 10.0001 | 0.020648 | 35.8034 | 0.98694 | 7.3334 | 1.0026 | -0.44655 | 7.3334 | -2.9855 | -26.6094 | -3.2748 |  | -1.4428 | -1.832 | -0.21743
2022-05-26 09:01 | 10.0001 | 0.020648 | 35.9539 | 0.98694 | 7.3752 | 1.0026 | -0.43523 | 7.3752 | -2.8349 | -22.7828 | -3.2099 |  | -1.3728 | -1.8371 | -0.12505
2022-05-26 10:01 | 10.0001 | 0.020648 | 36.1725 | 0.98694 | 7.417 | 1.0026 | -0.42554 | 7.417 | -2.6163 | -17.9087 | -3.1562 |  | -1.2704 | -1.8858 | -0.17982
2022-05-26 11:01 | 10.0001 | 0.0292 | 36.301 | 0.98694 | 7.4588 | 1.0025 | -0.41093 | 7.4588 | -2.4879 | -15.3985 | -3.065 |  | -1.2099 | -1.8551 | -0.16614
2022-05-26 12:00 | 10.0001 | 0.016859 | 36.4147 | 0.98694 | 7.4999 | 0.98711 | -0.40138 | 7.4999 | -2.3741 | -13.3819 | -3.0103 |  | -1.1561 | -1.8542 | -0.21066
2022-05-26 13:00 | 10.0001 | 0.011921 | 36.5549 | 0.98694 | 7.5417 | 1.0037 | -0.4013 | 7.5417 | -2.2339 | -11.1481 | -3.0265 |  | -1.0894 | -1.9371 | -0.26369
2022-05-26 14:01 | 10.0001 | 0.035763 | 36.7065 | 0.98694 | 7.5835 | 1.0025 | -0.38707 | 7.5835 | -2.0823 | -9.0286 | -2.9353 |  | -1.017 | -1.9184 | -0.15541
2022-05-26 15:01 | 10.0001 | 0.026656 | 36.8475 | 0.98694 | 7.6253 | 1.0026 | -0.37791 | 7.6253 | -1.9413 | -7.3163 | -2.8816 |  | -0.94936 | -1.9323 | -0.16568
2022-05-26 16:01 | 10.0001 | 0.011921 | 36.9729 | 0.98694 | 7.667 | 1.0026 | -0.36422 | 7.667 | -1.8159 | -5.9878 | -2.7925 |  | -0.88897 | -1.9035 | -0.15872
2022-05-26 17:00 | 10.0001 | 0 | 37.0147 | 0.98694 | 7.7081 | 0.9859 | -0.35767 | 7.7081 | -1.7741 | -5.5836 | -2.757 |  | -0.8688 | -1.8882 | -0.16422
2022-05-26 18:00 | 10.0001 | 0.0292 | 36.9743 | 0.98694 | 7.7499 | 1.0026 | -0.37209 | 7.7499 | -1.8145 | -5.9745 | -2.8837 |  | -0.88832 | -1.9953 | -0.21036
2022-05-26 19:00 | 10.0001 | 0.026656 | 36.8985 | 0.98694 | 7.7917 | 1.0026 | -0.37608 | 7.7917 | -1.8904 | -6.7551 | -2.9303 |  | -0.92485 | -2.0054 | -0.17826
2022-05-26 20:00 | 10.0001 | 0.020648 | 36.7664 | 0.98694 | 7.8334 | 1.0026 | -0.36916 | 7.8334 | -2.0224 | -8.2718 | -2.8918 |  | -0.98827 | -1.9035 | -0.13891
2022-05-26 21:01 | 10.0001 | 0.026656 | 36.7082 | 0.98694 | 7.8752 | 1.0025 | -0.37801 | 7.8752 | -2.0806 | -9.0072 | -2.9769 |  | -1.0162 | -1.9607 | -0.22787
2022-05-26 22:01 | 10.0001 | 0 | 36.5016 | 0.98694 | 7.917 | 1.0026 | -0.38945 | 7.917 | -2.2872 | -11.9654 | -3.0833 |  | -1.1148 | -1.9685 | -0.19729
2022-05-26 23:01 | 10.0001 | 0.020648 | 36.3977 | 0.98694 | 7.9588 | 1.0025 | -0.38995 | 7.9588 | -2.3911 | -13.671 | -3.1035 |  | -1.1641 | -1.9394 | -0.21477
2022-05-27 | 10.0001 | 0.035763 | 36.3329 | 0.98694 | 7.9998 | 0.98588 | -0.40264 | 7.9998 | -2.4559 | -14.8125 | -3.2211 |  | -1.1948 | -2.0263 | -0.21492
2022-05-27 01:00 | 10.0001 | 0.020648 | 36.1709 | 0.98694 | 8.0416 | 1.0025 | -0.40874 | 8.0416 | -2.6179 | -17.9411 | -3.2869 |  | -1.2712 | -2.0157 | -0.21072
2022-05-27 02:00 | 10.0001 | 0.0292 | 36.0707 | 0.98694 | 8.0834 | 1.0037 | -0.40888 | 8.0834 | -2.7182 | -20.083 | -3.3052 |  | -1.3183 | -1.9869 | -0.2387
2022-05-27 03:01 | 10.0001 | 0.020648 | 36.0163 | 0.98694 | 8.1252 | 1.0026 | -0.41102 | 8.1252 | -2.7725 | -21.3118 | -3.3396 |  | -1.3437 | -1.9959 | -0.22687
2022-05-27 04:01 | 10.0001 | 0.03154 | 35.969 | 0.98694 | 8.167 | 1.0026 | -0.41686 | 8.167 | -2.8198 | -22.4212 | -3.4045 |  | -1.3658 | -2.0387 | -0.23445
2022-05-27 05:01 | 10.0001 | 0.035763 | 35.8809 | 0.98694 | 8.2087 | 1.0026 | -0.41609 | 8.2087 | -2.9079 | -24.5896 | -3.4156 |  | -1.4068 | -2.0088 | -0.17193
2022-05-27 06:00 | 10.0001 | 0.0292 | 35.7697 | 0.98694 | 8.2498 | 0.9859 | -0.41488 | 8.2498 | -3.0191 | -27.5189 | -3.4227 |  | -1.4584 | -1.9643 | -0.17173
2022-05-27 07:00 | 10.0001 | 0.023842 | 35.8067 | 0.98694 | 8.2916 | 1.0026 | -0.42696 | 8.2916 | -2.9821 | -26.5193 | -3.5402 |  | -1.4412 | -2.099 | -0.21577
2022-05-27 08:00 | 10.0001 | NaN | 35.7835 | 0.98694 | 8.3334 | 1.0025 | -0.4258 | 8.3334 | -3.0053 | -27.1431 | -3.5483 |  | -1.452 | -2.0964 | -0.26441
2022-05-27 09:01 | 10.0001 | 0.020648 | 35.9459 | 0.98694 | 8.3751 | 1.0026 | -0.40516 | 8.3751 | -2.8429 | -22.9772 | -3.3933 |  | -1.3766 | -2.0168 | -0.17966
2022-05-27 10:01 | 10.0001 | 0.020648 | 36.1364 | 0.98694 | 8.4169 | 1.0026 | -0.3993 | 8.4169 | -2.6524 | -18.6605 | -3.3609 |  | -1.2874 | -2.0735 | -0.18771
2022-05-27 11:01 | 10.0001 | 0.03154 | 36.4515 | 0.98694 | 8.4587 | 1.0026 | -0.37457 | 8.4587 | -2.3374 | -12.7697 | -3.1684 |  | -1.1386 | -2.0298 | -0.17465
2022-05-27 12:00 | 10.0001 | 0.0292 | 36.7804 | 0.98694 | 8.4998 | 0.98591 | -0.36191 | 8.4998 | -2.0084 | -8.1017 | -3.0762 |  | -0.98158 | -2.0946 | -0.24036
2022-05-27 13:00 | 10.0001 | 0 | 36.9553 | 0.98694 | 8.5415 | 1.0026 | -0.34638 | 8.5415 | -1.8335 | -6.1641 | -2.9586 |  | -0.89748 | -2.0611 | -0.12404
2022-05-27 14:00 | 10.0001 | 0.020648 | 36.8926 | 0.98694 | 8.5833 | 1.0026 | -0.34741 | 8.5833 | -1.8962 | -6.8181 | -2.9819 |  | -0.92766 | -2.0543 | -0.13594
2022-05-27 15:00 | 10.0001 | 0.020648 | 36.8374 | 0.98694 | 8.6251 | 1.0026 | -0.35571 | 8.6251 | -1.9515 | -7.4316 | -3.0681 |  | -0.95423 | -2.1138 | -0.18159
2022-05-27 16:01 | 10.0001 | 0.026656 | 36.8895 | 0.98694 | 8.6669 | 1.0026 | -0.3492 | 8.6669 | -1.8994 | -6.8522 | -3.0265 |  | -0.92918 | -2.0973 | -0.19388
2022-05-27 17:01 | 10.0001 | 0.011921 | 36.996 | 0.98694 | 8.7086 | 1.0026 | -0.34858 | 8.7086 | -1.7929 | -5.7628 | -3.0356 |  | -0.87786 | -2.1578 | -0.26943
2022-05-27 18:01 | 10.0001 | 0.020648 | 37.1927 | 0.98694 | 8.7504 | 1.0026 | -0.33313 | 8.7504 | -1.5961 | -4.0662 | -2.9151 |  | -0.78272 | -2.1323 | -0.13694
2022-05-27 19:00 | 10.0001 | 0.0292 | 37.4258 | 0.98694 | 8.7915 | 0.98591 | -0.32063 | 8.7915 | -1.363 | -2.5323 | -2.8188 |  | -0.66947 | -2.1493 | -0.14394
2022-05-27 20:00 | 10.0001 | 0.023842 | 37.5984 | 0.98694 | 8.8333 | 1.0026 | -0.30443 | 8.8333 | -1.1905 | -1.6871 | -2.6891 |  | -0.58529 | -2.1038 | -0.20037
2022-05-27 21:00 | 10.0001 | 0.020648 | 37.5009 | 0.98694 | 8.8751 | 1.0026 | -0.31007 | 8.8751 | -1.288 | -2.1366 | -2.7519 |  | -0.63289 | -2.119 | -0.15838
2022-05-27 22:01 | 10.0001 | 0 | 37.2743 | 0.98694 | 8.9168 | 1.0026 | -0.32578 | 8.9168 | -1.5145 | -3.4739 | -2.9049 |  | -0.74313 | -2.1618 | -0.19334
2022-05-27 23:01 | 10.0001 | 0.020648 | 36.9616 | 0.98694 | 8.9586 | 1.0026 | -0.34292 | 8.9586 | -1.8272 | -6.1007 | -3.0721 |  | -0.89444 | -2.1777 | -0.23829
2022-05-28 00:01 | 10.0001 | 0.0292 | 36.7053 | 0.98694 | 9.0004 | 1.0025 | -0.35574 | 9.0004 | -2.0835 | -9.0444 | -3.2018 |  | -1.0175 | -2.1843 | -0.15789
2022-05-28 01:00 | 10.0001 | 0.033717 | 36.4529 | 0.98694 | 9.0415 | 0.98591 | -0.37385 | 9.0415 | -2.3359 | -12.7464 | -3.3801 |  | -1.1379 | -2.2422 | -0.22651
2022-05-28 02:00 | 10.0001 | 0.041295 | 36.2766 | 0.98694 | 9.0832 | 1.0025 | -0.37994 | 9.0832 | -2.5123 | -15.8562 | -3.4511 |  | -1.2214 | -2.2296 | -0.24278
2022-05-28 03:00 | 10.0001 | 0.033717 | 36.1869 | 0.98694 | 9.125 | 1.0025 | -0.38208 | 9.125 | -2.6019 | -17.6145 | -3.4865 |  | -1.2637 | -2.2229 | -0.22699
2022-05-28 04:00 | 10.0001 | 0.0292 | 35.9897 | 0.98694 | 9.1668 | 1.0025 | -0.38808 | 9.1668 | -2.7992 | -21.9324 | -3.5574 |  | -1.3561 | -2.2013 | -0.16267
2022-05-28 05:01 | 10.0001 | 0.026656 | 35.7545 | 0.98694 | 9.2085 | 1.0025 | -0.40173 | 9.2085 | -3.0343 | -27.9379 | -3.6993 |  | -1.4654 | -2.2339 | -0.22514
2022-05-28 06:01 | 10.0001 | NaN | 35.6083 | 0.98694 | 9.2503 | 1.0026 | -0.41032 | 9.2503 | -3.1806 | -32.1748 | -3.7956 |  | -1.5328 | -2.2627 | -0.29828
2022-05-28 07:01 | 10.0001 | 0 | 35.5843 | 0.98694 | 9.2921 | 1.0026 | -0.4052 | 9.2921 | -3.2045 | -32.907 | -3.7652 |  | -1.5438 | -2.2214 | -0.12231
2022-05-28 08:00 | 10.0001 | 0.026656 | 35.7701 | 0.98694 | 9.3332 | 0.98585 | -0.39799 | 9.3332 | -3.0187 | -27.5088 | -3.7145 |  | -1.4582 | -2.2563 | -0.15998
2022-05-28 09:00 | 10.0001 | 0.023842 | 35.9881 | 0.98694 | 9.3749 | 1.0025 | -0.39384 | 9.3749 | -2.8007 | -21.9696 | -3.6922 |  | -1.3569 | -2.3353 | -0.31865
2022-05-28 10:00 | 10.0001 | 0.039537 | 36.417 | 0.98694 | 9.4168 | 1.0037 | -0.36616 | 9.4168 | -2.3718 | -13.3428 | -3.448 |  | -1.155 | -2.2931 | -0.21963
2022-05-28 11:01 | 10.0001 | 0.026656 | 36.6826 | 0.98694 | 9.4585 | 1.0026 | -0.35019 | 9.4585 | -2.1063 | -9.344 | -3.3122 |  | -1.0284 | -2.2838 | -0.25409
2022-05-28 12:01 | 10.0001 | 0.016859 | 36.8244 | 0.98694 | 9.5003 | 1.0026 | -0.34395 | 9.5003 | -1.9644 | -7.5802 | -3.2677 |  | -0.96043 | -2.3072 | -0.21253
2022-05-28 13:01 | 10.0001 | 0.011921 | 36.807 | 0.98694 | 9.5421 | 1.0026 | -0.34425 | 9.5421 | -1.9819 | -7.7845 | -3.2849 |  | -0.96883 | -2.3161 | -0.25477
2022-05-28 14:00 | 10.0001 | 0.023842 | 36.7981 | 0.98694 | 9.5831 | 0.98583 | -0.33865 | 9.5831 | -1.9907 | -7.8893 | -3.2454 |  | -0.97308 | -2.2723 | -0.21806
2022-05-28 15:00 | 10.0001 | NaN | 36.8143 | 0.98694 | 9.6249 | 1.0026 | -0.33992 | 9.6249 | -1.9745 | -7.698 | -3.2717 |  | -0.9653 | -2.3064 | -0.19262
2022-05-28 16:00 | 10.0001 | 0.033717 | 36.8765 | 0.98694 | 9.6667 | 1.0026 | -0.33961 | 9.6667 | -1.9123 | -6.993 | -3.2829 |  | -0.9354 | -2.3475 | -0.25018
2022-05-28 17:01 | 10.0001 | 0.0292 | 36.8099 | 0.98694 | 9.7085 | 1.0026 | -0.33981 | 9.7085 | -1.9789 | -7.75 | -3.2991 |  | -0.96743 | -2.3316 | -0.1739
2022-05-28 18:01 | 10.0001 | NaN | 36.7916 | 0.98694 | 9.7502 | 1.0025 | -0.34293 | 9.7502 | -1.9972 | -7.9663 | -3.3437 |  | -0.97618 | -2.3675 | -0.23517
2022-05-28 19:01 | 10.0001 | 0.016859 | 36.7973 | 0.98694 | 9.792 | 1.0026 | -0.34416 | 9.792 | -1.9915 | -7.8983 | -3.37 |  | -0.97345 | -2.3966 | -0.24709
2022-05-28 20:00 | 10.0001 | 0.0292 | 36.6663 | 0.98694 | 9.8331 | 0.98588 | -0.3487 | 9.8331 | -2.1225 | -9.5624 | -3.4288 |  | -1.0362 | -2.3925 | -0.28878
2022-05-28 21:00 | 10.0001 | 0.020648 | 36.4701 | 0.98694 | 9.8749 | 1.0026 | -0.35112 | 9.8749 | -2.3187 | -12.466 | -3.4673 |  | -1.1297 | -2.3375 | -0.21854
2022-05-28 22:00 | 10.0001 | 0.03154 | 36.2207 | 0.98694 | 9.9166 | 1.0025 | -0.36047 | 9.9166 | -2.5681 | -16.9376 | -3.5747 |  | -1.2478 | -2.3269 | -0.16514
2022-05-28 23:00 | 10.0001 | 0.0292 | 35.9933 | 0.98694 | 9.9584 | 1.0026 | -0.37453 | 9.9584 | -2.7955 | -21.8463 | -3.7297 |  | -1.3544 | -2.3753 | -0.19764
2022-05-29 00:01 | 10.0001 | 0.011921 | 35.803 | 0.98694 | 10.0002 | 1.0026 | -0.38279 | 10.0002 | -2.9858 | -26.6181 | -3.828 |  | -1.4429 | -2.385 | -0.20082
2022-05-29 01:01 | 10.0001 | 0.016859 | 35.7056 | 0.98694 | 10.042 | 1.0026 | -0.38362 | 10.042 | -3.0832 | -29.3097 | -3.8523 |  | -1.488 | -2.3643 | -0.12204
2022-05-29 02:01 | 10.0001 | 0.023842 | 35.5306 | 0.98694 | 10.0837 | 1.0026 | -0.39077 | 10.0837 | -3.2582 | -34.5884 | -3.9404 |  | -1.5684 | -2.3721 | -0.14234
2022-05-29 03:00 | 10.0001 | 0.020648 | 35.398 | 0.98694 | 10.1248 | 0.98587 | -0.39429 | 10.1248 | -3.3908 | -38.9868 | -3.9921 |  | -1.6289 | -2.3632 | -0.14039
2022-05-29 04:00 | 10.0001 | 0.016859 | 35.2423 | 0.98694 | 10.1666 | 1.0026 | -0.40772 | 10.1666 | -3.5465 | -44.6072 | -4.1451 |  | -1.6994 | -2.4457 | -0.24444
2022-05-29 05:00 | 10.0001 | 0.020648 | 35.1111 | 0.98694 | 10.2084 | 1.0037 | -0.40764 | 10.2084 | -3.6777 | -49.7438 | -4.1613 |  | -1.7584 | -2.4029 | -0.16909
2022-05-29 06:01 | 10.0001 | 0.026656 | 35.0213 | 0.98694 | 10.2502 | 1.0026 | -0.41013 | 10.2502 | -3.7675 | -53.4764 | -4.2039 |  | -1.7985 | -2.4054 | -0.14266
2022-05-29 07:01 | 10.0001 | 0.023842 | 35.1564 | 0.98694 | 10.292 | 1.0026 | -0.40364 | 10.292 | -3.6324 | -47.9267 | -4.1542 |  | -1.7381 | -2.4162 | -0.19485
2022-05-29 08:01 | 10.0001 | 0.03154 | 35.306 | 0.98694 | 10.3337 | 1.0026 | -0.39642 | 10.3337 | -3.4829 | -42.2482 | -4.0965 |  | -1.6706 | -2.4258 | -0.16944
2022-05-29 09:00 | 10.0001 | 0.020648 | 35.3605 | 0.98694 | 10.3748 | 0.98586 | -0.39202 | 10.3748 | -3.4284 | -40.296 | -4.0671 |  | -1.646 | -2.4212 | -0.085828
2022-05-29 10:00 | 10.0001 | 0.026656 | 35.5357 | 0.98694 | 10.4166 | 1.0025 | -0.38578 | 10.4166 | -3.2531 | -34.426 | -4.0185 |  | -1.5661 | -2.4524 | -0.15937
2022-05-29 11:00 | 10.0001 | 0.0292 | 35.8025 | 0.98694 | 10.4584 | 1.0025 | -0.37242 | 10.4584 | -2.9863 | -26.6321 | -3.8949 |  | -1.4432 | -2.4517 | -0.16789
2022-05-29 12:01 | 10.0001 | 0 | 35.9414 | 0.98694 | 10.5001 | 1.0025 | -0.36823 | 10.5001 | -2.8474 | -23.0869 | -3.8665 |  | -1.3787 | -2.4878 | -0.18063
2022-05-29 13:01 | 10.0001 | 0 | 36.0444 | 0.98694 | 10.5419 | 1.0026 | -0.35918 | 10.5419 | -2.7444 | -20.6695 | -3.7864 |  | -1.3305 | -2.4559 | -0.13988
2022-05-29 14:01 | 10.0001 | 0.016859 | 36.0303 | 0.98694 | 10.5837 | 1.0025 | -0.36073 | 10.5837 | -2.7585 | -20.9902 | -3.8178 |  | -1.3371 | -2.4807 | -0.20832
2022-05-29 15:01 | 10.0001 | 0.026656 | 36.069 | 0.98694 | 10.6254 | 1.0025 | -0.35979 | 10.6254 | -2.7198 | -20.12 | -3.8229 |  | -1.319 | -2.5039 | -0.19736
2022-05-29 16:00 | 10.0001 | 0.020648 | 36.0866 | 0.98694 | 10.6665 | 0.98592 | -0.36087 | 10.6665 | -2.7022 | -19.7312 | -3.8493 |  | -1.3108 | -2.5385 | -0.19106
2022-05-29 17:00 | 10.0001 | 0.037697 | 36.0526 | 0.98694 | 10.7083 | 1.0026 | -0.35918 | 10.7083 | -2.7363 | -20.487 | -3.8462 |  | -1.3267 | -2.5195 | -0.18787
2022-05-29 18:00 | 10.0001 | 0.026656 | 36.1314 | 0.98694 | 10.7501 | 1.0026 | -0.35241 | 10.7501 | -2.6575 | -18.7671 | -3.7885 |  | -1.2898 | -2.4987 | -0.13124
2022-05-29 19:01 | 10.0001 | 0.023842 | 36.2657 | 0.98694 | 10.7918 | 1.0026 | -0.35152 | 10.7918 | -2.5232 | -16.0634 | -3.7935 |  | -1.2266 | -2.567 | -0.17044
2022-05-29 20:01 | 10.0001 | 0.026656 | 36.3132 | 0.98694 | 10.8336 | 1.0025 | -0.33922 | 10.8336 | -2.4757 | -15.1729 | -3.675 |  | -1.2041 | -2.4709 | -0.078278
2022-05-29 21:01 | 10.0001 | 0.020648 | 36.1401 | 0.98694 | 10.8754 | 1.0026 | -0.34761 | 10.8754 | -2.6487 | -18.5818 | -3.7804 |  | -1.2857 | -2.4947 | -0.15708
2022-05-29 22:00 | 10.0001 | 0.011921 | 35.9847 | 0.98694 | 10.9165 | 0.98588 | -0.354 | 10.9165 | -2.8041 | -22.0495 | -3.8645 |  | -1.3585 | -2.506 | -0.17912
2022-05-29 23:00 | 10.0001 | 0.020648 | 35.7112 | 0.98694 | 10.9582 | 1.0026 | -0.36319 | 10.9582 | -3.0777 | -29.1519 | -3.98 |  | -1.4854 | -2.4945 | -0.11928
2022-05-30 | 10.0001 | 0.023842 | 35.5979 | 0.98694 | 11 | 1.0026 | -0.37563 | 11 | -3.1909 | -32.4906 | -4.1319 |  | -1.5376 | -2.5944 | -0.20938
2022-05-30 01:01 | 10.0001 | 0.023842 | 35.4225 | 0.98694 | 11.0418 | 1.0025 | -0.37274 | 11.0418 | -3.3663 | -38.1465 | -4.1157 |  | -1.6177 | -2.498 | -0.13372
2022-05-30 02:01 | 10.0001 | 0.011921 | 35.3045 | 0.98694 | 11.0836 | 1.0026 | -0.38048 | 11.0836 | -3.4844 | -42.3031 | -4.2171 |  | -1.6713 | -2.5457 | -0.17371
2022-05-30 03:01 | 10.0001 | 0 | 35.2035 | 0.98694 | 11.1253 | 1.0026 | -0.38388 | 11.1253 | -3.5853 | -46.0877 | -4.2708 |  | -1.7169 | -2.5539 | -0.19055
2022-05-30 04:01 | 10.0001 | 0.011921 | 35.0963 | 0.98694 | 11.1671 | 1.0026 | -0.3858 | 11.1671 | -3.6925 | -50.3463 | -4.3083 |  | -1.765 | -2.5432 | -0.097481
2022-05-30 05:00 | 10.0001 | 0.026656 | 34.9975 | 0.98694 | 11.2082 | 0.98593 | -0.3918 | 11.2082 | -3.7913 | -54.4961 | -4.3913 |  | -1.8091 | -2.5822 | -0.17934
2022-05-30 06:00 | 10.0001 | 0.016859 | 34.8432 | 0.98694 | 11.25 | 1.0026 | -0.39917 | 11.25 | -3.9456 | -61.4259 | -4.4906 |  | -1.8775 | -2.6132 | -0.20782
2022-05-30 07:01 | 10.0001 | 0.023842 | 35.0019 | 0.98694 | 11.2918 | 1.0038 | -0.39222 | 11.2918 | -3.7869 | -54.3086 | -4.4288 |  | -1.8072 | -2.6217 | -0.20552
2022-05-30 08:01 | 10.0001 | NaN | 35.2064 | 0.98694 | 11.3336 | 1.0026 | -0.38389 | 11.3336 | -3.5824 | -45.9752 | -4.3508 |  | -1.7156 | -2.6352 | -0.20942
2022-05-30 09:01 | 10.0001 | 0.026656 | 35.2877 | 0.98694 | 11.3753 | 1.0026 | -0.38087 | 11.3753 | -3.5011 | -42.916 | -4.3326 |  | -1.6789 | -2.6537 | -0.23239
2022-05-30 10:00 | 10.0001 | 0.033717 | 35.3464 | 0.98694 | 11.4165 | 0.9872 | -0.37284 | 11.4165 | -3.4424 | -40.793 | -4.2566 |  | -1.6523 | -2.6043 | -0.15186
2022-05-30 11:00 | 10.0001 | 0 | 35.377 | 0.98694 | 11.4583 | 1.0026 | -0.37308 | 11.4583 | -3.4118 | -39.7141 | -4.2748 |  | -1.6384 | -2.6364 | -0.18475
2022-05-30 12:00 | 10.0001 | NaN | 35.592 | 0.98694 | 11.5 | 1.0026 | -0.36141 | 11.5 | -3.1969 | -32.672 | -4.1563 |  | -1.5403 | -2.616 | -0.12816
2022-05-30 13:01 | 10.0001 | 0.016859 | 35.7444 | 0.98694 | 11.5418 | 1.0036 | -0.35879 | 11.5418 | -3.0444 | -28.2171 | -4.1411 |  | -1.4701 | -2.671 | -0.21509
2022-05-30 14:01 | 10.0001 | NaN | 35.7156 | 0.98694 | 11.5836 | 1.0025 | -0.35679 | 11.5836 | -3.0732 | -29.0263 | -4.133 |  | -1.4834 | -2.6496 | -0.16886
2022-05-30 15:01 | 10.0001 | 0.020648 | 35.8203 | 0.98694 | 11.6254 | 1.0026 | -0.35481 | 11.6254 | -2.9685 | -26.1594 | -4.1249 |  | -1.435 | -2.6899 | -0.18603
2022-05-30 16:00 | 10.0001 | 0.026656 | 35.8807 | 0.98694 | 11.6665 | 0.98589 | -0.35105 | 11.6665 | -2.9081 | -24.594 | -4.0955 |  | -1.4069 | -2.6886 | -0.1501
2022-05-30 17:00 | 10.0001 | 0.03154 | 35.9388 | 0.98694 | 11.7082 | 1.0025 | -0.3472 | 11.7082 | -2.85 | -23.1492 | -4.0651 |  | -1.3799 | -2.6852 | -0.16575
2022-05-30 18:00 | 10.0001 | 0.026656 | 35.9989 | 0.98694 | 11.75 | 1.0026 | -0.34579 | 11.75 | -2.79 | -21.7165 | -4.063 |  | -1.3518 | -2.7112 | -0.21254
2022-05-30 19:01 | 10.0001 | 0.016859 | 36.1506 | 0.98694 | 11.7918 | 1.0025 | -0.33812 | 11.7918 | -2.6383 | -18.3632 | -3.9871 |  | -1.2808 | -2.7063 | -0.13934
2022-05-30 20:01 | 10.0001 | 0.020648 | 36.1496 | 0.98694 | 11.8336 | 1.0026 | -0.33821 | 11.8336 | -2.6392 | -18.3835 | -4.0023 |  | -1.2812 | -2.721 | -0.2502
2022-05-30 21:01 | 10.0001 | NaN | 36.1872 | 0.98694 | 11.8753 | 1.0025 | -0.33608 | 11.8753 | -2.6016 | -17.6088 | -3.9911 |  | -1.2635 | -2.7276 | -0.23289
2022-05-30 22:01 | 10.0001 | 0.016859 | 36.1211 | 0.98694 | 11.9171 | 1.0025 | -0.33975 | 11.9171 | -2.6678 | -18.9862 | -4.0489 |  | -1.2946 | -2.7543 | -0.24809
2022-05-30 23:00 | 10.0001 | 0.020648 | 36.1911 | 0.98694 | 11.9582 | 0.98594 | -0.33452 | 11.9582 | -2.5977 | -17.5296 | -4.0002 |  | -1.2617 | -2.7385 | -0.24402
2022-05-31 | 10.0001 | 0.020648 | 36.0895 | 0.98694 | 12 | 1.0026 | -0.33766 | 12 | -2.6994 | -19.6689 | -4.0519 |  | -1.3094 | -2.7425 | -0.1481
2022-05-31 01:00 | 10.0001 | 0 | 36.0459 | 0.98694 | 12.0417 | 1.0025 | -0.34196 | 12.0417 | -2.7429 | -20.6361 | -4.1178 |  | -1.3298 | -2.7879 | -0.28995
2022-05-31 02:01 | 10.0001 | 0.023842 | 35.9956 | 0.98694 | 12.0835 | 1.0026 | -0.34145 | 12.0835 | -2.7933 | -21.7938 | -4.1259 |  | -1.3534 | -2.7725 | -0.22677
2022-05-31 03:01 | 10.0001 | NaN | 35.9212 | 0.98694 | 12.1253 | 1.0026 | -0.34319 | 12.1253 | -2.8677 | -23.5819 | -4.1613 |  | -1.3881 | -2.7733 | -0.21941
2022-05-31 04:01 | 10.0001 | 0.020648 | 35.8294 | 0.98694 | 12.1671 | 1.0026 | -0.3416 | 12.1671 | -2.9594 | -25.9179 | -4.1563 |  | -1.4307 | -2.7256 | -0.18233
2022-05-31 05:00 | 10.0001 | 0.026656 | 35.7393 | 0.98694 | 12.2081 | 0.9858 | -0.35273 | 12.2081 | -3.0495 | -28.3595 | -4.3062 |  | -1.4724 | -2.8338 | -0.25156
2022-05-31 06:00 | 10.0001 | 0.016859 | 35.6328 | 0.98694 | 12.2499 | 1.0025 | -0.34955 | 12.2499 | -3.156 | -31.4361 | -4.2819 |  | -1.5215 | -2.7604 | -0.14721
2022-05-31 07:00 | 10.0001 | 0.026656 | 35.6223 | 0.98694 | 12.2917 | 1.0025 | -0.34622 | 12.2917 | -3.1666 | -31.7513 | -4.2556 |  | -1.5264 | -2.7292 | -0.10753
2022-05-31 08:01 | 10.0001 | 0 | 35.6261 | 0.98694 | 12.3334 | 1.0026 | -0.34899 | 12.3334 | -3.1627 | -31.6361 | -4.3042 |  | -1.5246 | -2.7796 | -0.14438
2022-05-31 09:01 | 10.0001 | 0.020648 | 35.7615 | 0.98694 | 12.3752 | 1.0025 | -0.3429 | 12.3752 | -3.0273 | -27.7441 | -4.2434 |  | -1.4622 | -2.7812 | -0.12759
2022-05-31 10:01 | 10.0001 | 0.011921 | 35.7372 | 0.98694 | 12.417 | 1.0026 | -0.34305 | 12.417 | -3.0516 | -28.4183 | -4.2596 |  | -1.4734 | -2.7862 | -0.18185
2022-05-31 11:01 | 10.0001 | 0.033717 | 36.2573 | 0.98694 | 12.4588 | 1.0025 | -0.32327 | 12.4588 | -2.5315 | -16.2233 | -4.0276 |  | -1.2305 | -2.7971 | -0.16061
2022-05-31 12:00 | 10.0001 | 0.03154 | 36.4568 | 0.98694 | 12.4998 | 0.98581 | -0.31743 | 12.4998 | -2.332 | -12.6819 | -3.9678 |  | -1.1361 | -2.8318 | -0.21582
2022-05-31 13:00 | 10.0001 | 0.0292 | 36.4794 | 0.98694 | 12.5416 | 1.0027 | -0.31298 | 12.5416 | -2.3094 | -12.3168 | -3.9252 |  | -1.1253 | -2.7999 | -0.12897
2022-05-31 14:00 | 10.0001 | 0.020648 | 36.8188 | 0.98694 | 12.5834 | 1.0025 | -0.3022 | 12.5834 | -1.97 | -7.6458 | -3.8026 |  | -0.96315 | -2.8395 | -0.18997
2022-05-31 15:01 | 10.0001 | 0.016859 | 37.1256 | 0.98694 | 12.6252 | 1.0025 | -0.28707 | 12.6252 | -1.6632 | -4.6012 | -3.6243 |  | -0.81523 | -2.8091 | -0.11921
2022-05-31 16:01 | 10.0001 | 0.016859 | 37.2293 | 0.98694 | 12.6669 | 1.0025 | -0.28436 | 12.6669 | -1.5595 | -3.7928 | -3.602 |  | -0.76497 | -2.8371 | -0.14841
2022-05-31 17:01 | 10.0001 | 0.023842 | 37.358 | 0.98694 | 12.7087 | 1.0025 | -0.27968 | 12.7087 | -1.4308 | -2.9294 | -3.5544 |  | -0.70247 | -2.8519 | -0.16664
2022-05-31 18:00 | 10.0001 | 0.016859 | 37.567 | 0.98694 | 12.7498 | 0.98592 | -0.27211 | 12.7498 | -1.2219 | -1.8242 | -3.4693 |  | -0.60063 | -2.8687 | -0.15749
2022-05-31 19:00 | 10.0001 | NaN | 37.5436 | 0.98694 | 12.7915 | 1.0025 | -0.27256 | 12.7915 | -1.2452 | -1.9308 | -3.4865 |  | -0.61203 | -2.8745 | -0.16823
2022-05-31 20:00 | 10.0001 | 0.016859 | 37.8109 | 0.98694 | 12.8333 | 1.0025 | -0.25944 | 12.8333 | -0.97797 | -0.93536 | -3.3295 |  | -0.48133 | -2.8481 | -0.12713
2022-05-31 21:00 | 10.0001 | 0.0292 | 37.5131 | 0.98694 | 12.8751 | 1.0025 | -0.26969 | 12.8751 | -1.2758 | -2.0764 | -3.4723 |  | -0.62694 | -2.8454 | -0.11785
2022-05-31 22:01 | 10.0001 | 0.026656 | 37.3941 | 0.98694 | 12.9169 | 1.0025 | -0.27486 | 12.9169 | -1.3947 | -2.7132 | -3.5504 |  | -0.68491 | -2.8654 | -0.11122
2022-05-31 23:01 | 10.0001 | 0.020648 | 37.2798 | 0.98694 | 12.9586 | 1.0025 | -0.28438 | 12.9586 | -1.509 | -3.436 | -3.6851 |  | -0.74045 | -2.9447 | -0.20604
2022-06-01 00:01 | 10.0001 | NaN | 37.0906 | 0.98694 | 13.0004 | 1.0025 | -0.28409 | 13.0004 | -1.6982 | -4.8977 | -3.6932 |  | -0.83216 | -2.8611 | -0.11854
2022-06-01 01:00 | 10.0001 | 0.016859 | 36.9971 | 0.98694 | 13.0415 | 0.98568 | -0.29018 | 13.0415 | -1.7917 | -5.7515 | -3.7844 |  | -0.8773 | -2.9071 | -0.11921
2022-06-01 02:00 | 10.0001 | 0.020648 | 36.9008 | 0.98694 | 13.0832 | 1.0025 | -0.29204 | 13.0832 | -1.8881 | -6.7306 | -3.8209 |  | -0.92374 | -2.8971 | -0.12469
2022-06-01 03:00 | 10.0001 | 0.033717 | 36.8013 | 0.98694 | 13.125 | 1.0025 | -0.30092 | 13.125 | -1.9875 | -7.8513 | -3.9496 |  | -0.97155 | -2.978 | -0.20482
2022-06-01 04:01 | 10.0001 | 0.023842 | 36.7372 | 0.98694 | 13.1668 | 1.0025 | -0.29635 | 13.1668 | -2.0516 | -8.6356 | -3.9019 |  | -1.0023 | -2.8997 | -0.17414
2022-06-01 05:01 | 10.0001 | 0.0292 | 36.6179 | 0.98694 | 13.2085 | 1.0022 | -0.29718 | 13.2085 | -2.1709 | -10.2311 | -3.9252 |  | -1.0593 | -2.8659 | -0.032117
2022-06-01 06:01 | 10.0001 | 0.0292 | 36.5409 | 0.98694 | 13.2503 | 1.0025 | -0.29616 | 13.2503 | -2.2479 | -11.359 | -3.9242 |  | -1.0961 | -2.8282 | -0.06778
2022-06-01 07:01 | 10.0001 | 0.0292 | 36.6265 | 0.98694 | 13.2921 | 1.0025 | -0.2982 | 13.2921 | -2.1623 | -10.1096 | -3.9638 |  | -1.0552 | -2.9085 | -0.17927
2022-06-01 08:00 | 10.0001 | NaN | 36.6968 | 0.98694 | 13.3332 | 0.98592 | -0.29584 | 13.3332 | -2.092 | -9.1558 | -3.9445 |  | -1.0216 | -2.9229 | -0.14332
2022-06-01 09:00 | 10.0001 | 0.011921 | 36.6571 | 0.98694 | 13.3749 | 1.0025 | -0.29537 | 13.3749 | -2.1317 | -9.6871 | -3.9506 |  | -1.0406 | -2.91 | -0.12877
2022-06-01 10:00 | 10.0001 | 0.037697 | 36.903 | 0.98694 | 13.4167 | 1.0026 | -0.28886 | 13.4167 | -1.8858 | -6.7065 | -3.8756 |  | -0.92266 | -2.9529 | -0.16679
2022-06-01 11:01 | 10.0001 | 0.023842 | 37.3985 | 0.98694 | 13.4585 | 1.0026 | -0.27012 | 13.4585 | -1.3903 | -2.6874 | -3.6355 |  | -0.68275 | -2.9527 | -0.15568
2022-06-01 12:01 | 10.0001 | 0.011921 | 37.5024 | 0.98694 | 13.5002 | 1.0026 | -0.26914 | 13.5002 | -1.2864 | -2.1289 | -3.6334 |  | -0.63214 | -3.0013 | -0.16948
2022-06-01 13:01 | 10.0001 | 0.011921 | 37.6752 | 0.98694 | 13.542 | 1.0025 | -0.2627 | 13.542 | -1.1136 | -1.3811 | -3.5574 |  | -0.54774 | -3.0097 | -0.20969
2022-06-01 14:00 | 10.0001 | 0.020648 | 37.7887 | 0.98694 | 13.5831 | 0.98585 | -0.25989 | 13.5831 | -1.0001 | -1.0004 | -3.5301 |  | -0.49218 | -3.0379 | -0.19846
2022-06-01 15:00 | 10.0001 | 0.020648 | 37.9621 | 0.98694 | 13.6249 | 1.0026 | -0.24853 | 13.6249 | -0.82673 | -0.56506 | -3.3862 |  | -0.40714 | -2.9791 | -0.17003
2022-06-01 16:00 | 10.0001 | 0.020648 | 38.1042 | 0.98694 | 13.6666 | 1.0026 | -0.24547 | 13.6666 | -0.68463 | -0.3209 | -3.3548 |  | -0.33733 | -3.0175 | -0.18046
2022-06-01 17:00 | 10.0001 | 0.033717 | 38.2858 | 0.98694 | 13.7084 | 1.0025 | -0.23807 | 13.7084 | -0.50306 | -0.12731 | -3.2636 |  | -0.24798 | -3.0156 | -0.16374
2022-06-01 18:01 | 10.0001 | 0.023842 | 38.4882 | 0.98694 | 13.7502 | 1.0025 | -0.23521 | 13.7502 | -0.3006 | -0.027163 | -3.2342 |  | -0.14824 | -3.086 | -0.21723
2022-06-01 19:01 | 10.0001 | 0.0292 | 38.6389 | 0.98694 | 13.792 | 1.0024 | -0.22657 | 13.792 | -0.14988 | -0.0033669 | -3.1248 |  | -0.073923 | -3.0509 | -0.17631
2022-06-01 20:01 | 10.0001 | NaN | 38.7127 | 0.98694 | 13.8337 | 1.0025 | -0.22398 | 13.8337 | -0.07616 | -0.00044176 | -3.0985 |  | -0.037565 | -3.0609 | -0.21266
2022-06-01 21:00 | 10.0001 | 0.0292 | 38.5015 | 0.98694 | 13.8748 | 0.9858 | -0.23157 | 13.8748 | -0.28732 | -0.02372 | -3.2129 |  | -0.14169 | -3.0713 | -0.22592
2022-06-01 22:00 | 10.0001 | 0.03154 | 38.1686 | 0.98694 | 13.9166 | 1.0025 | -0.24289 | 13.9166 | -0.62023 | -0.23859 | -3.3801 |  | -0.30565 | -3.0745 | -0.20909
2022-06-01 23:00 | 10.0001 | 0.033717 | 37.9527 | 0.98694 | 13.9583 | 1.0025 | -0.25305 | 13.9583 | -0.83612 | -0.58453 | -3.5321 |  | -0.41175 | -3.1204 | -0.17575
2022-06-02 00:01 | 10.0001 | 0.041295 | 37.7716 | 0.98694 | 14.0001 | 1.0025 | -0.25794 | 14.0001 | -1.0172 | -1.0524 | -3.6111 |  | -0.50053 | -3.1106 | -0.24962
2022-06-02 01:01 | 10.0001 | 0 | 37.6137 | 0.98694 | 14.0419 | 1.0025 | -0.26388 | 14.0419 | -1.1751 | -1.6227 | -3.7054 |  | -0.57779 | -3.1276 | -0.22038
2022-06-02 02:01 | 10.0001 | 0.0292 | 37.5805 | 0.98694 | 14.0836 | 1.0025 | -0.26677 | 14.0836 | -1.2084 | -1.7644 | -3.7571 |  | -0.59404 | -3.163 | -0.26576
2022-06-02 03:01 | 10.0001 | 0.011921 | 37.4652 | 0.98694 | 14.1254 | 1.0025 | -0.26562 | 14.1254 | -1.3237 | -2.3192 | -3.752 |  | -0.65029 | -3.1017 | -0.12363
2022-06-02 04:00 | 10.0001 | 0.020648 | 37.401 | 0.98694 | 14.1665 | 0.98592 | -0.27014 | 14.1665 | -1.3878 | -2.6731 | -3.827 |  | -0.68155 | -3.1454 | -0.24583
2022-06-02 05:00 | 10.0001 | 0.020648 | 37.231 | 0.98694 | 14.2083 | 1.0025 | -0.27241 | 14.2083 | -1.5578 | -3.7803 | -3.8705 |  | -0.76414 | -3.1064 | -0.24055
2022-06-02 06:00 | 10.0001 | 0 | 37.1237 | 0.98694 | 14.25 | 1.0026 | -0.27652 | 14.25 | -1.6651 | -4.6169 | -3.9404 |  | -0.81615 | -3.1243 | -0.2962
2022-06-02 07:01 | 10.0001 | 0.026656 | 37.2439 | 0.98694 | 14.2918 | 1.0025 | -0.27238 | 14.2918 | -1.545 | -3.6876 | -3.8928 |  | -0.75791 | -3.1349 | -0.22644
2022-06-02 08:01 | 10.0001 | 0.011921 | 37.3198 | 0.98694 | 14.3336 | 1.0026 | -0.27138 | 14.3336 | -1.4691 | -3.1705 | -3.8898 |  | -0.72105 | -3.1687 | -0.24575
2022-06-02 09:01 | 10.0001 | 0.023842 | 37.4782 | 0.98694 | 14.3754 | 1.0025 | -0.26305 | 14.3754 | -1.3106 | -2.2512 | -3.7814 |  | -0.64392 | -3.1374 | -0.22739
2022-06-02 10:00 | 10.0001 | 0 | 37.6086 | 0.98694 | 14.4164 | 0.9859 | -0.26201 | 14.4164 | -1.1802 | -1.6439 | -3.7773 |  | -0.58029 | -3.197 | -0.24415
2022-06-02 11:00 | 10.0001 | 0.026656 | 37.6109 | 0.98694 | 14.4582 | 1.0025 | -0.25677 | 14.4582 | -1.1779 | -1.6342 | -3.7125 |  | -0.57915 | -3.1333 | -0.18065
2022-06-02 12:00 | 10.0001 | 0.011921 | 37.647 | 0.98694 | 14.5 | 1.0037 | -0.25366 | 14.5 | -1.1418 | -1.4885 | -3.678 |  | -0.56151 | -3.1165 | -0.11524
2022-06-02 13:01 | 10.0001 | 0.023842 | 37.6922 | 0.98694 | 14.5418 | 1.0025 | -0.25335 | 14.5418 | -1.0967 | -1.3189 | -3.6841 |  | -0.53944 | -3.1447 | -0.13498
2022-06-02 14:01 | 10.0001 | 0.0292 | 37.734 | 0.98694 | 14.5836 | 1.0025 | -0.25137 | 14.5836 | -1.0548 | -1.1735 | -3.6659 |  | -0.51895 | -3.1469 | -0.10896
2022-06-02 15:01 | 10.0001 | 0.023842 | 37.7656 | 0.98694 | 14.6254 | 1.0025 | -0.25315 | 14.6254 | -1.0232 | -1.0712 | -3.7023 |  | -0.50347 | -3.1989 | -0.21969
2022-06-02 16:00 | 10.0001 | 0.016859 | 37.7605 | 0.98694 | 14.6664 | 0.98583 | -0.25264 | 14.6664 | -1.0284 | -1.0875 | -3.7054 |  | -0.50601 | -3.1994 | -0.18193
2022-06-02 17:00 | 10.0001 | 0.020648 | 37.8123 | 0.98694 | 14.7082 | 1.0025 | -0.24834 | 14.7082 | -0.97657 | -0.93134 | -3.6527 |  | -0.48064 | -3.1721 | -0.15646
2022-06-02 18:00 | 10.0001 | 0.020648 | 37.8279 | 0.98694 | 14.75 | 1.0025 | -0.24881 | 14.75 | -0.96088 | -0.88717 | -3.6699 |  | -0.47295 | -3.197 | -0.111
2022-06-02 19:00 | 10.0001 | NaN | 37.8688 | 0.98694 | 14.7917 | 1.0026 | -0.24783 | 14.7917 | -0.92004 | -0.77879 | -3.6659 |  | -0.45293 | -3.2129 | -0.16209
2022-06-02 20:01 | 10.0001 | 0.0292 | 37.9688 | 0.98694 | 14.8335 | 1.0025 | -0.24508 | 14.8335 | -0.81998 | -0.55133 | -3.6355 |  | -0.40383 | -3.2316 | -0.17078
2022-06-02 21:01 | 10.0001 | 0.03154 | 37.944 | 0.98694 | 14.8753 | 1.0025 | -0.2461 | 14.8753 | -0.84483 | -0.60299 | -3.6608 |  | -0.41603 | -3.2448 | -0.17343
2022-06-02 22:01 | 10.0001 | 0.033717 | 37.8747 | 0.98694 | 14.9171 | 1.0025 | -0.24976 | 14.9171 | -0.91412 | -0.76384 | -3.7256 |  | -0.45002 | -3.2756 | -0.20104
2022-06-02 23:00 | 10.0001 | NaN | 37.7071 | 0.98694 | 14.9582 | 0.98707 | -0.25117 | 14.9582 | -1.0817 | -1.2658 | -3.7571 |  | -0.53214 | -3.2249 | -0.10457
2022-06-03 | 10.0001 | 0.020648 | 37.5359 | 0.98694 | 15 | 1.0025 | -0.25966 | 15 | -1.253 | -1.967 | -3.8949 |  | -0.61581 | -3.279 | -0.16846
2022-06-03 01:01 | 10.0001 | 0.011921 | 37.4993 | 0.98694 | 15.0418 | 1.0036 | -0.25732 | 15.0418 | -1.2895 | -2.1442 | -3.8705 |  | -0.63364 | -3.2369 | -0.10932
2022-06-03 02:01 | 10.0001 | 0 | 37.4333 | 0.98694 | 15.0835 | 1.0025 | -0.25849 | 15.0835 | -1.3556 | -2.4909 | -3.8989 |  | -0.66583 | -3.2331 | -0.070067
2022-06-03 03:01 | 10.0001 | 0.03154 | 37.3091 | 0.98694 | 15.1253 | 1.0026 | -0.26045 | 15.1253 | -1.4797 | -3.24 | -3.9394 |  | -0.72623 | -3.2132 | -0.11151
2022-06-03 04:01 | 10.0001 | 0 | 37.2463 | 0.98694 | 15.1671 | 1.0025 | -0.26455 | 15.1671 | -1.5425 | -3.6699 | -4.0124 |  | -0.75671 | -3.2557 | -0.1102
2022-06-03 05:00 | 10.0001 | 0.03154 | 37.1969 | 0.98694 | 15.2082 | 0.98579 | -0.26823 | 15.2082 | -1.5919 | -4.0345 | -4.0793 |  | -0.7807 | -3.2986 | -0.19218
2022-06-03 06:00 | 10.0001 | 0.023842 | 37.1602 | 0.98694 | 15.25 | 1.0036 | -0.27135 | 15.25 | -1.6287 | -4.32 | -4.138 |  | -0.79849 | -3.3395 | -0.21525
2022-06-03 07:01 | 10.0001 | 0.011921 | 37.1756 | 0.98694 | 15.2918 | 1.0037 | -0.26875 | 15.2918 | -1.6133 | -4.1987 | -4.1097 |  | -0.79103 | -3.3186 | -0.18371
2022-06-03 08:01 | 10.0001 | 0.033717 | 37.2574 | 0.98694 | 15.3336 | 1.0026 | -0.26412 | 15.3336 | -1.5314 | -3.5917 | -4.0499 |  | -0.75135 | -3.2985 | -0.1298
2022-06-03 09:01 | 10.0001 | 0.016859 | 37.4881 | 0.98694 | 15.3753 | 1.0026 | -0.25945 | 15.3753 | -1.3008 | -2.2009 | -3.9891 |  | -0.63913 | -3.35 | -0.21251
2022-06-03 10:00 | 10.0001 | 0.0292 | 37.6175 | 0.98694 | 15.4164 | 0.98586 | -0.25337 | 15.4164 | -1.1714 | -1.6073 | -3.906 |  | -0.57597 | -3.33 | -0.133
2022-06-03 11:00 | 10.0001 | 0.023842 | 37.7923 | 0.98694 | 15.4582 | 1.0025 | -0.24862 | 15.4582 | -0.99654 | -0.98964 | -3.8432 |  | -0.49042 | -3.3528 | -0.21944
2022-06-03 12:00 | 10.0001 | 0.020648 | 37.8637 | 0.98694 | 15.5 | 1.0025 | -0.24468 | 15.5 | -0.92515 | -0.79185 | -3.7925 |  | -0.45544 | -3.3371 | -0.22058
2022-06-03 13:00 | 10.0001 | 0.0292 | 37.9743 | 0.98694 | 15.5417 | 1.0025 | -0.24089 | 15.5417 | -0.81451 | -0.54036 | -3.7439 |  | -0.40114 | -3.3427 | -0.1981
2022-06-03 14:01 | 10.0001 | 0.020648 | 38.1178 | 0.98694 | 15.5835 | 1.0025 | -0.23719 | 15.5835 | -0.67099 | -0.3021 | -3.6963 |  | -0.33062 | -3.3656 | -0.21874
2022-06-03 15:01 | 10.0001 | 0.039537 | 38.1767 | 0.98694 | 15.6253 | 1.0025 | -0.23481 | 15.6253 | -0.61213 | -0.22937 | -3.6689 |  | -0.30167 | -3.3672 | -0.16838
2022-06-03 16:01 | 10.0001 | 0.020648 | 38.2238 | 0.98694 | 15.6671 | 1.0026 | -0.23593 | 15.6671 | -0.565 | -0.18036 | -3.6963 |  | -0.27848 | -3.4178 | -0.21828
2022-06-03 17:00 | 10.0001 | 0.016859 | 38.2686 | 0.98694 | 15.7081 | 0.98585 | -0.23518 | 15.7081 | -0.52026 | -0.14082 | -3.6942 |  | -0.25645 | -3.4378 | -0.26575
2022-06-03 18:00 | 10.0001 | 0.023842 | 38.2435 | 0.98694 | 15.7499 | 1.0026 | -0.23179 | 15.7499 | -0.54533 | -0.16217 | -3.6507 |  | -0.2688 | -3.3819 | -0.18491
2022-06-03 19:00 | 10.0001 | 0.011921 | 38.2613 | 0.98694 | 15.7917 | 1.0025 | -0.23342 | 15.7917 | -0.52754 | -0.14681 | -3.6861 |  | -0.26004 | -3.4261 | -0.21317
2022-06-03 20:01 | 10.0001 | NaN | 38.2024 | 0.98694 | 15.8334 | 1.0025 | -0.23505 | 15.8334 | -0.58646 | -0.20171 | -3.7216 |  | -0.28904 | -3.4326 | -0.20093
2022-06-03 21:01 | 10.0001 | 0.016859 | 38.1202 | 0.98694 | 15.8752 | 1.0025 | -0.24068 | 15.8752 | -0.66859 | -0.29887 | -3.8209 |  | -0.32944 | -3.4914 | -0.24669
2022-06-03 22:01 | 10.0001 | 0.016859 | 37.9324 | 0.98694 | 15.917 | 1.0026 | -0.24438 | 15.917 | -0.85639 | -0.62809 | -3.8898 |  | -0.4217 | -3.4681 | -0.19247
2022-06-03 23:01 | 10.0001 | 0.016859 | 37.9133 | 0.98694 | 15.9588 | 1.0025 | -0.24698 | 15.9588 | -0.87554 | -0.67116 | -3.9415 |  | -0.4311 | -3.5104 | -0.28528
2022-06-04 | 10.0001 | 0.016859 | 37.7118 | 0.98694 | 15.9998 | 0.98586 | -0.25217 | 15.9998 | -1.077 | -1.2493 | -4.0347 |  | -0.52983 | -3.5048 | -0.22583
2022-06-04 01:00 | 10.0001 | 0.011921 | 37.5754 | 0.98694 | 16.0416 | 1.0025 | -0.25417 | 16.0416 | -1.2134 | -1.7866 | -4.0772 |  | -0.5965 | -3.4807 | -0.24387
2022-06-04 02:00 | 10.0001 | 0.026656 | 37.5024 | 0.98694 | 16.0834 | 1.0024 | -0.26113 | 16.0834 | -1.2864 | -2.129 | -4.1998 |  | -0.63214 | -3.5677 | -0.33467
2022-06-04 03:01 | 10.0001 | 0.020648 | 37.3138 | 0.98694 | 16.1252 | 1.0025 | -0.26139 | 16.1252 | -1.475 | -3.2091 | -4.215 |  | -0.72394 | -3.4911 | -0.27794
2022-06-04 04:01 | 10.0001 | 0 | 37.2755 | 0.98694 | 16.1669 | 1.0026 | -0.26028 | 16.1669 | -1.5134 | -3.466 | -4.2079 |  | -0.74257 | -3.4654 | -0.20972
2022-06-04 05:01 | 10.0001 | 0.020648 | 37.1992 | 0.98694 | 16.2087 | 1.0025 | -0.26161 | 16.2087 | -1.5896 | -4.0169 | -4.2404 |  | -0.77958 | -3.4608 | -0.16214
2022-06-04 06:00 | 10.0001 | 0.016859 | 37.0216 | 0.98694 | 16.2498 | 0.98583 | -0.26594 | 16.2498 | -1.7672 | -5.5193 | -4.3214 |  | -0.8655 | -3.4559 | -0.11641
2022-06-04 07:00 | 10.0001 | 0.016859 | 36.9525 | 0.98694 | 16.2915 | 1.0025 | -0.26718 | 16.2915 | -1.8364 | -6.1927 | -4.3528 |  | -0.89884 | -3.454 | -0.1354
2022-06-04 08:00 | 10.0001 | 0.020648 | 36.879 | 0.98694 | 16.3333 | 1.0025 | -0.26954 | 16.3333 | -1.9098 | -6.9654 | -4.4025 |  | -0.93419 | -3.4683 | -0.16981
2022-06-04 09:00 | 10.0001 | 0.020648 | 36.8596 | 0.98694 | 16.3751 | 1.0025 | -0.2704 | 16.3751 | -1.9292 | -7.1804 | -4.4278 |  | -0.94354 | -3.4843 | -0.13436
2022-06-04 10:01 | 10.0001 | 0 | 36.7915 | 0.98694 | 16.4169 | 1.0025 | -0.27329 | 16.4169 | -1.9973 | -7.9677 | -4.4866 |  | -0.97624 | -3.5103 | -0.18023
2022-06-04 11:01 | 10.0001 | 0.020648 | 36.8447 | 0.98694 | 16.4586 | 1.0026 | -0.27315 | 16.4586 | -1.9441 | -7.3476 | -4.4957 |  | -0.95068 | -3.545 | -0.19217
2022-06-04 12:01 | 10.0001 | 0.03154 | 36.9333 | 0.98694 | 16.5004 | 1.0025 | -0.26804 | 16.5004 | -1.8555 | -6.3884 | -4.4227 |  | -0.90807 | -3.5147 | -0.17752
2022-06-04 13:00 | 10.0001 | NaN | 36.9937 | 0.98694 | 16.5415 | 0.98585 | -0.26694 | 16.5415 | -1.7951 | -5.7843 | -4.4157 |  | -0.87894 | -3.5367 | -0.19403
2022-06-04 14:00 | 10.0001 | 0.03154 | 37.1319 | 0.98694 | 16.5833 | 1.0025 | -0.26053 | 16.5833 | -1.657 | -4.5492 | -4.3204 |  | -0.81219 | -3.5082 | -0.14262
2022-06-04 15:00 | 10.0001 | 0.026656 | 37.2908 | 0.98694 | 16.625 | 1.0025 | -0.25768 | 16.625 | -1.4981 | -3.3619 | -4.2839 |  | -0.73514 | -3.5488 | -0.18161
2022-06-04 16:01 | 10.0001 | 0.016859 | 37.3556 | 0.98694 | 16.6668 | 1.0025 | -0.25618 | 16.6668 | -1.4332 | -2.9439 | -4.2697 |  | -0.70362 | -3.5661 | -0.14839
2022-06-04 17:01 | 10.0001 | 0.023842 | 37.4307 | 0.98694 | 16.7086 | 1.0025 | -0.25099 | 16.7086 | -1.3581 | -2.5051 | -4.1938 |  | -0.66708 | -3.5267 | -0.088855
2022-06-04 18:01 | 10.0001 | 0.016859 | 37.525 | 0.98694 | 16.7503 | 1.0025 | -0.24892 | 16.7503 | -1.2639 | -2.0188 | -4.1694 |  | -0.62112 | -3.5483 | -0.16638
2022-06-04 19:01 | 10.0001 | 0.033717 | 37.6295 | 0.98694 | 16.7921 | 1.0025 | -0.24365 | 16.7921 | -1.1593 | -1.558 | -4.0914 |  | -0.57006 | -3.5214 | -0.095226
2022-06-04 20:00 | 10.0001 | 0.016859 | 37.6336 | 0.98694 | 16.8332 | 0.98588 | -0.24613 | 16.8332 | -1.1552 | -1.5417 | -4.1431 |  | -0.56808 | -3.575 | -0.1425
2022-06-04 21:00 | 10.0001 | 0.026656 | 37.5825 | 0.98694 | 16.875 | 1.0026 | -0.24816 | 16.875 | -1.2063 | -1.7553 | -4.1877 |  | -0.59302 | -3.5947 | -0.10324
2022-06-04 22:00 | 10.0001 | 0.020648 | 37.48 | 0.98694 | 16.9167 | 1.0025 | -0.25042 | 16.9167 | -1.3088 | -2.242 | -4.2363 |  | -0.64305 | -3.5933 | -0.12521
2022-06-04 23:01 | 10.0001 | 0.020648 | 37.3574 | 0.98694 | 16.9585 | 1.0025 | -0.25088 | 16.9585 | -1.4314 | -2.9328 | -4.2545 |  | -0.70274 | -3.5518 | -0.041458
2022-06-05 00:01 | 10.0001 | 0.03154 | 37.1882 | 0.98694 | 17.0003 | 1.0025 | -0.25646 | 17.0003 | -1.6007 | -4.1012 | -4.3599 |  | -0.78493 | -3.575 | -0.070118
2022-06-05 01:01 | 10.0001 | 0.0292 | 37.0532 | 0.98694 | 17.042 | 1.0026 | -0.25952 | 17.042 | -1.7356 | -5.2284 | -4.4227 |  | -0.85023 | -3.5725 | -0.091738
2022-06-05 02:00 | 10.0001 | 0.016859 | 37.0282 | 0.98694 | 17.0831 | 0.98584 | -0.25812 | 17.0831 | -1.7606 | -5.4576 | -4.4096 |  | -0.86231 | -3.5473 | 0.020428
2022-06-05 03:00 | 10.0001 | 0.011921 | 36.9615 | 0.98694 | 17.1249 | 1.0025 | -0.25974 | 17.1249 | -1.8274 | -6.1021 | -4.4481 |  | -0.8945 | -3.5536 | -0.062498
2022-06-05 04:00 | 10.0001 | 0.033717 | 36.931 | 0.98694 | 17.1667 | 1.0025 | -0.26224 | 17.1667 | -1.8578 | -6.4124 | -4.5018 |  | -0.90919 | -3.5926 | -0.12725
2022-06-05 05:00 | 10.0001 | 0.026656 | 36.934 | 0.98694 | 17.2084 | 1.0025 | -0.26249 | 17.2084 | -1.8548 | -6.3814 | -4.517 |  | -0.90774 | -3.6092 | -0.14849
2022-06-05 06:01 | 10.0001 | 0.020648 | 36.8774 | 0.98694 | 17.2502 | 1.0025 | -0.26332 | 17.2502 | -1.9114 | -6.9836 | -4.5423 |  | -0.93499 | -3.6073 | -0.15133
2022-06-05 07:01 | 10.0001 | 0 | 36.8257 | 0.98694 | 17.292 | 1.0025 | -0.26632 | 17.292 | -1.9632 | -7.566 | -4.6051 |  | -0.95985 | -3.6453 | -0.19121
2022-06-05 08:01 | 10.0001 | 0.016859 | 36.6998 | 0.98694 | 17.3337 | 1.0025 | -0.26591 | 17.3337 | -2.089 | -9.1167 | -4.6092 |  | -1.0202 | -3.589 | -0.12064
2022-06-05 09:00 | 10.0001 | 0.016859 | 36.7289 | 0.98694 | 17.3748 | 0.98584 | -0.26615 | 17.3748 | -2.0599 | -8.7407 | -4.6244 |  | -1.0063 | -3.6181 | -0.13388
2022-06-05 10:00 | 10.0001 | 0.020648 | 36.7054 | 0.98694 | 17.4166 | 1.0026 | -0.26755 | 17.4166 | -2.0834 | -9.0432 | -4.6598 |  | -1.0175 | -3.6423 | -0.13203
2022-06-05 11:00 | 10.0001 | 0.011921 | 36.7393 | 0.98694 | 17.4584 | 1.0025 | -0.26604 | 17.4584 | -2.0496 | -8.6095 | -4.6446 |  | -1.0013 | -3.6433 | -0.09836
2022-06-05 12:01 | 10.0001 | 0.011921 | 36.7599 | 0.98694 | 17.5001 | 1.0025 | -0.26656 | 17.5001 | -2.0289 | -8.352 | -4.6649 |  | -0.9914 | -3.6735 | -0.15887
2022-06-05 13:01 | 10.0001 | 0.020648 | 36.7465 | 0.98694 | 17.5419 | 1.0026 | -0.26229 | 17.5419 | -2.0423 | -8.519 | -4.6011 |  | -0.99784 | -3.6032 | -0.066487
2022-06-05 14:01 | 10.0001 | 0.020648 | 36.847 | 0.98694 | 17.5837 | 1.0025 | -0.26098 | 17.5837 | -1.9418 | -7.3222 | -4.5889 |  | -0.9496 | -3.6393 | -0.13104
2022-06-05 15:00 | 10.0001 | 0.0292 | 36.8436 | 0.98694 | 17.6248 | 0.98585 | -0.26157 | 17.6248 | -1.9452 | -7.3604 | -4.6102 |  | -0.95122 | -3.659 | -0.1102
2022-06-05 16:00 | 10.0001 | 0 | 36.8432 | 0.98694 | 17.6665 | 1.0025 | -0.26119 | 17.6665 | -1.9457 | -7.3656 | -4.6142 |  | -0.95144 | -3.6628 | -0.096696
2022-06-05 17:00 | 10.0001 | 0.023842 | 36.8764 | 0.98694 | 17.7083 | 1.0025 | -0.25948 | 17.7083 | -1.9124 | -6.9943 | -4.595 |  | -0.93545 | -3.6595 | -0.1329
2022-06-05 18:00 | 10.0001 | 0.016859 | 36.8656 | 0.98694 | 17.7501 | 1.0025 | -0.25613 | 17.7501 | -1.9232 | -7.1136 | -4.5464 |  | -0.94065 | -3.6057 | -0.057406
2022-06-05 19:01 | 10.0001 | NaN | 36.8616 | 0.98694 | 17.7919 | 1.0025 | -0.25895 | 17.7919 | -1.9272 | -7.1578 | -4.6072 |  | -0.94257 | -3.6646 | -0.14327
2022-06-05 20:01 | 10.0001 | NaN | 36.846 | 0.98694 | 17.8336 | 1.0025 | -0.25982 | 17.8336 | -1.9428 | -7.3332 | -4.6335 |  | -0.95007 | -3.6834 | -0.10837
2022-06-05 21:01 | 10.0001 | 0 | 36.8875 | 0.98694 | 17.8754 | 1.0025 | -0.26091 | 17.8754 | -1.9013 | -6.8731 | -4.6639 |  | -0.93011 | -3.7338 | -0.13907
2022-06-05 22:00 | 10.0001 | 0.037697 | 36.9478 | 0.98694 | 17.9165 | 0.98586 | -0.2602 | 17.9165 | -1.841 | -6.24 | -4.6619 |  | -0.90109 | -3.7608 | -0.16756
2022-06-05 23:00 | 10.0001 | 0.03154 | 36.9273 | 0.98694 | 17.9583 | 1.0038 | -0.25514 | 17.9583 | -1.8615 | -6.4507 | -4.5818 |  | -0.91096 | -3.6709 | -0.11907
2022-06-06 | 10.0001 | 0.011921 | 36.8833 | 0.98694 | 18.0001 | 1.0025 | -0.25775 | 18.0001 | -1.9055 | -6.9188 | -4.6396 |  | -0.93213 | -3.7074 | -0.13247
2022-06-06 01:01 | 10.0001 | 0.020648 | 36.8477 | 0.98694 | 18.0418 | 1.0025 | -0.25609 | 18.0418 | -1.9412 | -7.3145 | -4.6203 |  | -0.94928 | -3.671 | -0.098557
2022-06-06 02:01 | 10.0001 | 0 | 36.8933 | 0.98694 | 18.0837 | 1.0036 | -0.25477 | 18.0837 | -1.8955 | -6.8107 | -4.6072 |  | -0.92733 | -3.6798 | -0.13248
2022-06-06 03:01 | 10.0001 | 0.023842 | 36.7684 | 0.98694 | 18.1254 | 1.0025 | -0.26106 | 18.1254 | -2.0204 | -8.2477 | -4.7318 |  | -0.98733 | -3.7444 | -0.19077
2022-06-06 04:00 | 10.0001 | 0 | 36.7108 | 0.98694 | 18.1665 | 0.98591 | -0.26242 | 18.1665 | -2.078 | -8.9731 | -4.7672 |  | -1.0149 | -3.7523 | -0.15976
2022-06-06 05:00 | 10.0001 | 0.011921 | 36.6653 | 0.98694 | 18.2083 | 1.0025 | -0.26076 | 18.2083 | -2.1236 | -9.5764 | -4.748 |  | -1.0367 | -3.7113 | -0.10205
2022-06-06 06:00 | 10.0001 | 0.020648 | 36.659 | 0.98694 | 18.2501 | 1.0026 | -0.26183 | 18.2501 | -2.1299 | -9.6619 | -4.7784 |  | -1.0397 | -3.7387 | -0.13136
2022-06-06 07:01 | 10.0001 | 0.020648 | 36.6253 | 0.98694 | 18.2918 | 1.0025 | -0.26295 | 18.2918 | -2.1635 | -10.1269 | -4.8098 |  | -1.0558 | -3.754 | -0.10873
2022-06-06 08:01 | 10.0001 | 0 | 36.6078 | 0.98694 | 18.3336 | 1.0025 | -0.26274 | 18.3336 | -2.1811 | -10.3754 | -4.8169 |  | -1.0642 | -3.7527 | -0.16375
2022-06-06 09:01 | 10.0001 | 0.03154 | 36.5184 | 0.98694 | 18.3754 | 1.0025 | -0.26368 | 18.3754 | -2.2705 | -11.7043 | -4.8453 |  | -1.1068 | -3.7385 | -0.12028
2022-06-06 10:00 | 10.0001 | 0.016859 | 36.521 | 0.98694 | 18.4164 | 0.98589 | -0.26342 | 18.4164 | -2.2678 | -11.6638 | -4.8513 |  | -1.1055 | -3.7458 | -0.10347
2022-06-06 11:00 | 10.0001 | 0.0292 | 36.4931 | 0.98694 | 18.4582 | 1.0025 | -0.2664 | 18.4582 | -2.2957 | -12.0987 | -4.9172 |  | -1.1188 | -3.7984 | -0.15508
2022-06-06 12:00 | 10.0001 | 0.020648 | 36.5171 | 0.98694 | 18.5 | 1.0025 | -0.26344 | 18.5 | -2.2717 | -11.723 | -4.8736 |  | -1.1074 | -3.7663 | -0.092767
2022-06-06 13:00 | 10.0001 | 0.03154 | 36.6118 | 0.98694 | 18.5418 | 1.0025 | -0.26219 | 18.5418 | -2.177 | -10.3178 | -4.8615 |  | -1.0623 | -3.7992 | -0.19602
2022-06-06 14:01 | 10.0001 | 0.020648 | 36.721 | 0.98694 | 18.5835 | 1.0025 | -0.26182 | 18.5835 | -2.0678 | -8.8416 | -4.8655 |  | -1.01 | -3.8555 | -0.21621
2022-06-06 15:01 | 10.0001 | 0.016859 | 36.7647 | 0.98694 | 18.6253 | 1.0025 | -0.2565 | 18.6253 | -2.0241 | -8.2929 | -4.7774 |  | -0.9891 | -3.7883 | -0.12924
2022-06-06 16:01 | 10.0001 | 0.023842 | 36.8683 | 0.98694 | 18.6671 | 1.0025 | -0.25229 | 18.6671 | -1.9205 | -7.0837 | -4.7095 |  | -0.93936 | -3.7701 | -0.10727
2022-06-06 17:00 | 10.0001 | 0.051962 | 36.9477 | 0.98694 | 18.7082 | 0.98589 | -0.25314 | 18.7082 | -1.8411 | -6.2407 | -4.7358 |  | -0.90113 | -3.8347 | -0.17519
2022-06-06 18:00 | 10.0001 | 0.020648 | 36.9905 | 0.98694 | 18.7499 | 1.0025 | -0.25171 | 18.7499 | -1.7983 | -5.8157 | -4.7196 |  | -0.8805 | -3.8391 | -0.23345
2022-06-06 19:00 | 10.0001 | 0.03154 | 36.9729 | 0.98694 | 18.7917 | 1.0025 | -0.25056 | 18.7917 | -1.8159 | -5.9878 | -4.7085 |  | -0.88897 | -3.8195 | -0.15494
2022-06-06 20:01 | 10.0001 | 0.016859 | 37.0488 | 0.98694 | 18.8335 | 1.0026 | -0.25076 | 18.8335 | -1.74 | -5.268 | -4.7227 |  | -0.85234 | -3.8703 | -0.18692
2022-06-06 21:01 | 10.0001 | 0.023842 | 37.0781 | 0.98694 | 18.8752 | 1.0025 | -0.25192 | 18.8752 | -1.7107 | -5.0065 | -4.7551 |  | -0.8382 | -3.9169 | -0.18313
2022-06-06 22:01 | 10.0001 | 0.033717 | 37.0623 | 0.98694 | 18.917 | 1.0025 | -0.24869 | 18.917 | -1.7266 | -5.1469 | -4.7044 |  | -0.84585 | -3.8586 | -0.097744
2022-06-06 23:00 | 10.0001 | 0.0292 | 37.029 | 0.98694 | 18.9581 | 0.98587 | -0.2489 | 18.9581 | -1.7598 | -5.4501 | -4.7186 |  | -0.86192 | -3.8567 | -0.18587
2022-06-07 | 10.0001 | 0.011921 | 37.0972 | 0.98694 | 18.9999 | 1.0026 | -0.24958 | 18.9999 | -1.6917 | -4.8411 | -4.7419 |  | -0.82898 | -3.9129 | -0.20553
2022-06-07 01:00 | 10.0001 | 0.020648 | 36.9731 | 0.98694 | 19.0416 | 1.0025 | -0.25158 | 19.0416 | -1.8157 | -5.9863 | -4.7905 |  | -0.8889 | -3.9016 | -0.23065
2022-06-07 02:00 | 10.0001 | NaN | 36.9361 | 0.98694 | 19.0834 | 1.0025 | -0.24981 | 19.0834 | -1.8527 | -6.3599 | -4.7672 |  | -0.90673 | -3.8605 | -0.18073
2022-06-07 03:01 | 10.0001 | 0.016859 | 36.8314 | 0.98694 | 19.1252 | 1.0026 | -0.25218 | 19.1252 | -1.9575 | -7.5004 | -4.823 |  | -0.95711 | -3.8659 | -0.12144
2022-06-07 04:01 | 10.0001 | 0.020648 | 36.7246 | 0.98694 | 19.167 | 1.0026 | -0.25612 | 19.167 | -2.0642 | -8.7954 | -4.9091 |  | -1.0083 | -3.9008 | -0.1484
2022-06-07 05:01 | 10.0001 | 0.026656 | 36.7815 | 0.98694 | 19.2087 | 1.0025 | -0.25472 | 19.2087 | -2.0073 | -8.0879 | -4.8929 |  | -0.98103 | -3.9118 | -0.20049
2022-06-07 06:00 | 10.0001 | 0 | 36.785 | 0.98694 | 19.2498 | 0.98578 | -0.25465 | 19.2498 | -2.0038 | -8.0463 | -4.902 |  | -0.97938 | -3.9226 | -0.18401
2022-06-07 07:00 | 10.0001 | 0.051962 | 36.8918 | 0.98694 | 19.2916 | 1.0025 | -0.25037 | 19.2916 | -1.897 | -6.8269 | -4.8301 |  | -0.92806 | -3.902 | -0.14805
2022-06-07 08:00 | 10.0001 | 0.020648 | 36.8796 | 0.98694 | 19.3333 | 1.0026 | -0.24904 | 19.3333 | -1.9092 | -6.959 | -4.8149 |  | -0.9339 | -3.881 | -0.12829
2022-06-07 09:01 | 10.0001 | 0.0292 | 36.5089 | 0.98694 | 19.3751 | 1.0025 | -0.2609 | 19.3751 | -2.2799 | -11.8513 | -5.055 |  | -1.1113 | -3.9437 | -0.20528
2022-06-07 10:01 | 10.0001 | 0.026656 | 36.5247 | 0.98694 | 19.4169 | 1.0026 | -0.25956 | 19.4169 | -2.2641 | -11.606 | -5.0398 |  | -1.1038 | -3.936 | -0.19016
2022-06-07 11:01 | 10.0001 | 0.020648 | 36.835 | 0.98694 | 19.4587 | 1.0025 | -0.25504 | 19.4587 | -1.9538 | -7.4581 | -4.9628 |  | -0.95534 | -4.0075 | -0.20896
2022-06-07 12:01 | 10.0001 | 0.011921 | 36.9639 | 0.98694 | 19.5004 | 1.0025 | -0.24883 | 19.5004 | -1.8249 | -6.0774 | -4.8524 |  | -0.89332 | -3.959 | -0.19269
2022-06-07 13:00 | 10.0001 | 0.059605 | 37.0852 | 0.98694 | 19.5415 | 0.98592 | -0.24779 | 19.5415 | -1.7037 | -4.9449 | -4.8422 |  | -0.83479 | -4.0074 | -0.20827
2022-06-07 14:00 | 10.0001 | 0.011921 | 37.2986 | 0.98694 | 19.5833 | 1.0025 | -0.24219 | 19.5833 | -1.4903 | -3.3097 | -4.7429 |  | -0.73135 | -4.0116 | -0.15612
2022-06-07 15:00 | 10.0001 | 0.016859 | 37.4502 | 0.98694 | 19.6251 | 1.0026 | -0.236 | 19.6251 | -1.3386 | -2.3988 | -4.6315 |  | -0.65758 | -3.9739 | -0.18566
2022-06-07 16:01 | 10.0001 | 0.016859 | 37.5364 | 0.98694 | 19.6668 | 1.0025 | -0.23488 | 19.6668 | -1.2525 | -1.9647 | -4.6193 |  | -0.61557 | -4.0037 | -0.23367
2022-06-07 17:01 | 10.0001 | 0.033717 | 37.5909 | 0.98694 | 19.7086 | 1.0025 | -0.23397 | 19.7086 | -1.198 | -1.7192 | -4.6112 |  | -0.58896 | -4.0222 | -0.18746
2022-06-07 18:01 | 10.0001 | 0.026656 | 37.6075 | 0.98694 | 19.7504 | 1.0026 | -0.23301 | 19.7504 | -1.1813 | -1.6485 | -4.6021 |  | -0.58082 | -4.0213 | -0.18207
2022-06-07 19:00 | 10.0001 | 0.026656 | 37.5782 | 0.98694 | 19.7914 | 0.98581 | -0.23381 | 19.7914 | -1.2106 | -1.7744 | -4.6274 |  | -0.59515 | -4.0323 | -0.21281
2022-06-07 20:00 | 10.0001 | 0.023842 | 37.64 | 0.98694 | 19.8332 | 1.0025 | -0.23158 | 19.8332 | -1.1489 | -1.5164 | -4.593 |  | -0.56497 | -4.028 | -0.15772
2022-06-07 21:00 | 10.0001 | 0.016859 | 37.6147 | 0.98694 | 19.875 | 1.0025 | -0.23339 | 19.875 | -1.1741 | -1.6185 | -4.6386 |  | -0.5773 | -4.0613 | -0.14441
2022-06-07 22:00 | 10.0001 | 0 | 37.5535 | 0.98694 | 19.9168 | 1.0025 | -0.23224 | 19.9168 | -1.2353 | -1.8852 | -4.6254 |  | -0.60721 | -4.0182 | -0.15965
2022-06-07 23:01 | 10.0001 | 0.035763 | 37.4755 | 0.98694 | 19.9585 | 1.0026 | -0.23246 | 19.9585 | -1.3133 | -2.2653 | -4.6396 |  | -0.64525 | -3.9943 | -0.13757
2022-06-08 00:01 | 10.0001 | 0.0292 | 37.4017 | 0.98694 | 20.0003 | 1.0025 | -0.2337 | 20.0003 | -1.3871 | -2.669 | -4.674 |  | -0.6812 | -3.9928 | -0.07986
2022-06-08 01:01 | 10.0001 | 0.039537 | 37.3969 | 0.98694 | 20.0421 | 1.0025 | -0.23458 | 20.0421 | -1.3919 | -2.6968 | -4.7014 |  | -0.68353 | -4.0178 | -0.11615
2022-06-08 02:00 | 10.0001 | 0 | 37.3992 | 0.98694 | 20.0831 | 0.98583 | -0.23329 | 20.0831 | -1.3896 | -2.6833 | -4.6852 |  | -0.68241 | -4.0028 | -0.14229
2022-06-08 03:00 | 10.0001 | 0.011921 | 37.3624 | 0.98694 | 20.1249 | 1.0026 | -0.23442 | 20.1249 | -1.4265 | -2.9026 | -4.7176 |  | -0.70034 | -4.0173 | -0.15143
2022-06-08 04:00 | 10.0001 | 0.011921 | 37.3079 | 0.98694 | 20.1667 | 1.0026 | -0.23654 | 20.1667 | -1.4809 | -3.2476 | -4.7703 |  | -0.7268 | -4.0435 | -0.14274
2022-06-08 05:01 | 10.0001 | 0.0292 | 37.2226 | 0.98694 | 20.2085 | 1.0026 | -0.23666 | 20.2085 | -1.5662 | -3.842 | -4.7824 |  | -0.76823 | -4.0142 | -0.10239
2022-06-08 06:01 | 10.0001 | 0.026656 | 37.2456 | 0.98694 | 20.2502 | 1.0025 | -0.23687 | 20.2502 | -1.5433 | -3.6755 | -4.7966 |  | -0.75708 | -4.0395 | -0.11687
2022-06-08 07:01 | 10.0001 | 0.026656 | 37.2042 | 0.98694 | 20.292 | 1.0026 | -0.23653 | 20.292 | -1.5846 | -3.979 | -4.7997 |  | -0.77715 | -4.0225 | -0.12046
2022-06-08 08:01 | 10.0001 | 0.03154 | 37.2677 | 0.98694 | 20.3338 | 1.0025 | -0.23684 | 20.3338 | -1.5211 | -3.5195 | -4.8159 |  | -0.74634 | -4.0695 | -0.1885
2022-06-08 09:00 | 10.0001 | 0.016859 | 37.3672 | 0.98694 | 20.3749 | 0.98582 | -0.23368 | 20.3749 | -1.4217 | -2.8733 | -4.7612 |  | -0.698 | -4.0632 | -0.11949
2022-06-08 10:00 | 10.0001 | 0.03154 | 37.3395 | 0.98694 | 20.4166 | 1.0026 | -0.23528 | 20.4166 | -1.4493 | -3.0441 | -4.8037 |  | -0.71144 | -4.0923 | -0.15628
2022-06-08 11:00 | 10.0001 | 0.020648 | 37.4139 | 0.98694 | 20.4584 | 1.0026 | -0.23243 | 20.4584 | -1.375 | -2.5995 | -4.7551 |  | -0.67528 | -4.0798 | -0.072366
2022-06-08 12:01 | 10.0001 | 0.023842 | 37.1511 | 0.98694 | 20.5002 | 1.0026 | -0.24001 | 20.5002 | -1.6378 | -4.393 | -4.9202 |  | -0.8029 | -4.1173 | -0.15834
2022-06-08 13:01 | 10.0001 | 0.016859 | 37.4562 | 0.98694 | 20.5419 | 1.0026 | -0.23271 | 20.5419 | -1.3326 | -2.3667 | -4.7804 |  | -0.65466 | -4.1258 | -0.11827
2022-06-08 14:01 | 10.0001 | NaN | 37.8206 | 0.98694 | 20.5837 | 1.0025 | -0.22643 | 20.5837 | -0.96827 | -0.90781 | -4.6609 |  | -0.47657 | -4.1843 | -0.17263
2022-06-08 15:00 | 10.0001 | 0.020648 | 37.9524 | 0.98694 | 20.6248 | 0.98594 | -0.2225 | 20.6248 | -0.83638 | -0.58508 | -4.5889 |  | -0.41188 | -4.177 | -0.2032
2022-06-08 16:00 | 10.0001 | 0.033717 | 38.0017 | 0.98694 | 20.6666 | 1.0025 | -0.21935 | 20.6666 | -0.78714 | -0.48769 | -4.5332 |  | -0.3877 | -4.1455 | -0.14178
2022-06-08 17:00 | 10.0001 | 0.023842 | 38.2907 | 0.98694 | 20.7083 | 1.0025 | -0.21401 | 20.7083 | -0.49814 | -0.12361 | -4.4319 |  | -0.24557 | -4.1863 | -0.16409
2022-06-08 18:01 | 10.0001 | 0 | 38.4403 | 0.98694 | 20.7501 | 1.0025 | -0.21051 | 20.7501 | -0.34854 | -0.04234 | -4.368 |  | -0.17187 | -4.1962 | -0.17494
2022-06-08 19:01 | 10.0001 | 0.016859 | 38.553 | 0.98694 | 20.7919 | 1.0025 | -0.20696 | 20.7919 | -0.23587 | -0.013123 | -4.3032 |  | -0.11633 | -4.1869 | -0.15453
2022-06-08 20:01 | 10.0001 | 0.037697 | 38.8623 | 0.98694 | 20.8337 | 1.0026 | -0.19755 | 20.8337 | 0.073469 | 0.00039656 | -4.1157 |  | 0.036237 | -4.152 | -0.12392
2022-06-08 21:01 | 10.0001 | 0.03154 | 38.5386 | 0.98694 | 20.8754 | 1.0025 | -0.20648 | 20.8754 | -0.2502 | -0.015663 | -4.3103 |  | -0.12339 | -4.1869 | -0.12557
2022-06-08 22:00 | 10.0001 | 0.011921 | 38.107 | 0.98694 | 20.9165 | 0.98591 | -0.21532 | 20.9165 | -0.68183 | -0.31698 | -4.5038 |  | -0.33595 | -4.1679 | -0.14971
2022-06-08 23:00 | 10.0001 | 0.026656 | 38.0318 | 0.98694 | 20.9583 | 1.0025 | -0.22074 | 20.9583 | -0.75705 | -0.43388 | -4.6264 |  | -0.37292 | -4.2535 | -0.25923
2022-06-09 | 10.0001 | 0.023842 | 37.9965 | 0.98694 | 21 | 1.0025 | -0.21915 | 21 | -0.79232 | -0.4974 | -4.6021 |  | -0.39025 | -4.2118 | -0.21907
2022-06-09 01:01 | 10.0001 | 0.016859 | 37.8518 | 0.98694 | 21.0418 | 1.0026 | -0.22348 | 21.0418 | -0.93705 | -0.82278 | -4.7024 |  | -0.46127 | -4.2411 | -0.22334
2022-06-09 02:01 | 10.0001 | 0.020648 | 37.7808 | 0.98694 | 21.0836 | 1.0026 | -0.22525 | 21.0836 | -1.0081 | -1.0244 | -4.749 |  | -0.49607 | -4.2529 | -0.25006
2022-06-09 03:01 | 10.0001 | 0.033717 | 37.714 | 0.98694 | 21.1254 | 1.0026 | -0.22399 | 21.1254 | -1.0748 | -1.2416 | -4.7318 |  | -0.52874 | -4.203 | -0.18571
2022-06-09 04:00 | 10.0001 | 0.023842 | 37.6874 | 0.98694 | 21.1664 | 0.98588 | -0.22575 | 21.1664 | -1.1014 | -1.3362 | -4.7784 |  | -0.54178 | -4.2366 | -0.19317
2022-06-09 05:00 | 10.0001 | 0.016859 | 37.6704 | 0.98694 | 21.2082 | 1.0026 | -0.22483 | 21.2082 | -1.1184 | -1.399 | -4.7683 |  | -0.55008 | -4.2182 | -0.20401
2022-06-09 06:00 | 10.0001 | 0.03154 | 37.6459 | 0.98694 | 21.25 | 1.0025 | -0.2231 | 21.25 | -1.1429 | -1.4928 | -4.7409 |  | -0.56204 | -4.1789 | -0.13935
2022-06-09 07:00 | 10.0001 | 0.033717 | 37.7588 | 0.98694 | 21.2918 | 1.0025 | -0.221 | 21.2918 | -1.03 | -1.0926 | -4.7054 |  | -0.5068 | -4.1986 | -0.17616
2022-06-09 08:01 | 10.0001 | 0.026656 | 37.9641 | 0.98694 | 21.3335 | 1.0025 | -0.21914 | 21.3335 | -0.82468 | -0.56086 | -4.675 |  | -0.40614 | -4.2689 | -0.19941
2022-06-09 09:01 | 10.0001 | NaN | 38.0685 | 0.98694 | 21.3753 | 1.0026 | -0.21369 | 21.3753 | -0.72034 | -0.37378 | -4.5676 |  | -0.35488 | -4.2128 | -0.14953
2022-06-09 10:01 | 10.0001 | 0.020648 | 38.2095 | 0.98694 | 21.4171 | 1.0026 | -0.21346 | 21.4171 | -0.57929 | -0.19439 | -4.5717 |  | -0.28551 | -4.2862 | -0.19381
2022-06-09 11:00 | 10.0001 | 0.023842 | 38.2996 | 0.98694 | 21.4582 | 0.98588 | -0.20852 | 21.4582 | -0.48921 | -0.11708 | -4.4744 |  | -0.24117 | -4.2333 | -0.15349
2022-06-09 12:00 | 10.0001 | 0.023842 | 38.3509 | 0.98694 | 21.4999 | 1.0025 | -0.20835 | 21.4999 | -0.43787 | -0.083956 | -4.4795 |  | -0.21588 | -4.2636 | -0.1463
2022-06-09 13:00 | 10.0001 | 0.026656 | 38.6881 | 0.98694 | 21.5417 | 1.0025 | -0.20249 | 21.5417 | -0.10076 | -0.0010228 | -4.362 |  | -0.049695 | -4.3123 | -0.18655
2022-06-09 14:01 | 10.0001 | 0.016859 | 38.9225 | 0.98694 | 21.5835 | 1.0025 | -0.19787 | 21.5835 | 0.13367 | 0.0023883 | -4.2708 |  | 0.065928 | -4.3367 | -0.15245
2022-06-09 15:01 | 10.0001 | 0.020648 | 38.9695 | 0.98694 | 21.6252 | 1.0025 | -0.19543 | 21.6252 | 0.18066 | 0.0058964 | -4.2262 |  | 0.089102 | -4.3153 | -0.13819
2022-06-09 16:01 | 10.0001 | 0 | 38.6933 | 0.98694 | 21.667 | 1.0026 | -0.20174 | 21.667 | -0.095504 | -0.0008711 | -4.3711 |  | -0.047105 | -4.324 | -0.1784
2022-06-09 17:00 | 10.0001 | 0.023842 | 38.3994 | 0.98694 | 21.7081 | 0.98594 | -0.20654 | 21.7081 | -0.38945 | -0.059069 | -4.4835 |  | -0.19203 | -4.2915 | -0.10524
2022-06-09 18:00 | 10.0001 | 0.020648 | 38.1977 | 0.98694 | 21.7499 | 1.0036 | -0.21145 | 21.7499 | -0.59115 | -0.20659 | -4.599 |  | -0.29135 | -4.3077 | -0.11155
2022-06-09 19:00 | 10.0001 | 0.020648 | 38.0464 | 0.98694 | 21.7917 | 1.0036 | -0.21402 | 21.7917 | -0.74241 | -0.40919 | -4.6639 |  | -0.36573 | -4.2982 | -0.11132
2022-06-09 20:01 | 10.0001 | 0.0292 | 37.8344 | 0.98694 | 21.8335 | 1.0026 | -0.21667 | 21.8335 | -0.95439 | -0.86931 | -4.7308 |  | -0.46977 | -4.261 | -0.10904
2022-06-09 21:01 | 10.0001 | 0 | 37.697 | 0.98694 | 21.8753 | 1.0025 | -0.22261 | 21.8753 | -1.0918 | -1.3016 | -4.8696 |  | -0.53708 | -4.3325 | -0.14564
2022-06-09 22:01 | 10.0001 | 0.020648 | 37.5518 | 0.98694 | 21.9171 | 1.0026 | -0.22486 | 21.9171 | -1.237 | -1.8929 | -4.9283 |  | -0.60803 | -4.3203 | -0.15238
2022-06-09 23:00 | 10.0001 | 0.03154 | 37.4242 | 0.98694 | 21.9581 | 0.98592 | -0.22647 | 21.9581 | -1.3646 | -2.5412 | -4.9729 |  | -0.67024 | -4.3027 | -0.049208
2022-06-10 | 10.0001 | 0.020648 | 37.2981 | 0.98694 | 21.9999 | 1.0026 | -0.22941 | 21.9999 | -1.4907 | -3.3128 | -5.0469 |  | -0.73158 | -4.3153 | -0.10349
2022-06-10 01:00 | 10.0001 | 0.016859 | 37.176 | 0.98694 | 22.0417 | 1.0026 | -0.23394 | 22.0417 | -1.6129 | -4.1956 | -5.1563 |  | -0.79083 | -4.3655 | -0.12438
2022-06-10 02:01 | 10.0001 | 0.011921 | 37.0187 | 0.98694 | 22.0835 | 1.0026 | -0.23542 | 22.0835 | -1.7701 | -5.5464 | -5.1989 |  | -0.86689 | -4.332 | -0.07906
2022-06-10 03:01 | 10.0001 | 0.020648 | 36.9899 | 0.98694 | 22.1252 | 1.0026 | -0.23639 | 22.1252 | -1.7989 | -5.8217 | -5.2303 |  | -0.8808 | -4.3495 | -0.14647
2022-06-10 04:01 | 10.0001 | 0.03154 | 36.8976 | 0.98694 | 22.1671 | 1.0039 | -0.23746 | 22.1671 | -1.8913 | -6.7649 | -5.2637 |  | -0.92529 | -4.3384 | -0.10177
2022-06-10 05:00 | 10.0001 | 0.026656 | 36.8406 | 0.98694 | 22.2081 | 0.98594 | -0.24062 | 22.2081 | -1.9482 | -7.3949 | -5.3438 |  | -0.95268 | -4.3911 | -0.17293
2022-06-10 06:00 | 10.0001 | 0.016859 | 36.8218 | 0.98694 | 22.25 | 1.0038 | -0.2439 | 22.25 | -1.967 | -7.6107 | -5.4269 |  | -0.9617 | -4.4652 | -0.2863
2022-06-10 07:00 | 10.0001 | 0.037697 | 36.8329 | 0.98694 | 22.2917 | 1.0026 | -0.24008 | 22.2917 | -1.9559 | -7.4824 | -5.3519 |  | -0.95636 | -4.3955 | -0.19689
2022-06-10 08:01 | 10.0001 | 0.023842 | 37.0987 | 0.98694 | 22.3336 | 1.0038 | -0.23433 | 22.3336 | -1.6901 | -4.8277 | -5.2333 |  | -0.82822 | -4.4051 | -0.1362
2022-06-10 09:01 | 10.0001 | 0.016859 | 37.4065 | 0.98694 | 22.3753 | 1.0026 | -0.22696 | 22.3753 | -1.3823 | -2.6414 | -5.0783 |  | -0.67887 | -4.3994 | -0.18668
2022-06-10 10:01 | 10.0001 | 0 | 37.7529 | 0.98694 | 22.4171 | 1.0026 | -0.22048 | 22.4171 | -1.0359 | -1.1117 | -4.9425 |  | -0.50971 | -4.4328 | -0.14664
2022-06-10 11:00 | 10.0001 | NaN | 38.0053 | 0.98694 | 22.4582 | 0.98593 | -0.21317 | 22.4582 | -0.78348 | -0.48092 | -4.7875 |  | -0.3859 | -4.4016 | -0.16835
2022-06-10 12:00 | 10.0001 | 0.016859 | 38.1395 | 0.98694 | 22.5 | 1.0026 | -0.20877 | 22.5 | -0.64933 | -0.27378 | -4.6973 |  | -0.31997 | -4.3774 | -0.11375
2022-06-10 13:00 | 10.0001 | 0.023842 | 38.2901 | 0.98694 | 22.5417 | 1.0026 | -0.20555 | 22.5417 | -0.49868 | -0.12401 | -4.6335 |  | -0.24583 | -4.3877 | -0.075408
2022-06-10 14:01 | 10.0001 | 0.023842 | 38.3792 | 0.98694 | 22.5835 | 1.0026 | -0.2062 | 22.5835 | -0.40966 | -0.068748 | -4.6568 |  | -0.20198 | -4.4548 | -0.11813
2022-06-10 15:01 | 10.0001 | 0.020648 | 38.427 | 0.98694 | 22.6253 | 1.0026 | -0.20255 | 22.6253 | -0.36184 | -0.047377 | -4.5828 |  | -0.17842 | -4.4044 | -0.089128
2022-06-10 16:01 | 10.0001 | 0.016859 | 38.4274 | 0.98694 | 22.6671 | 1.0026 | -0.20415 | 22.6671 | -0.36147 | -0.047229 | -4.6274 |  | -0.17824 | -4.4492 | -0.12521
2022-06-10 17:00 | 10.0001 | 0.026656 | 38.5391 | 0.98694 | 22.7081 | 0.98588 | -0.19887 | 22.7081 | -0.24977 | -0.015581 | -4.516 |  | -0.12318 | -4.3928 | -0.10127
2022-06-10 18:00 | 10.0001 | 0.023842 | 38.673 | 0.98694 | 22.7499 | 1.0025 | -0.1981 | 22.7499 | -0.11585 | -0.001555 | -4.5068 |  | -0.057142 | -4.4497 | -0.142
2022-06-10 19:00 | 10.0001 | 0.0292 | 38.8808 | 0.98694 | 22.7917 | 1.0025 | -0.19405 | 22.7917 | 0.092001 | 0.00077872 | -4.4227 |  | 0.045378 | -4.4681 | -0.16996
2022-06-10 20:01 | 10.0001 | 0.0292 | 39.0253 | 0.98694 | 22.8335 | 1.0025 | -0.19201 | 22.8335 | 0.23649 | 0.013227 | -4.3842 |  | 0.11663 | -4.5009 | -0.23989
2022-06-10 21:01 | 10.0001 | 0.026656 | 39.018 | 0.98694 | 22.8752 | 1.0026 | -0.19139 | 22.8752 | 0.22914 | 0.012031 | -4.3782 |  | 0.11301 | -4.4912 | -0.15868
2022-06-10 22:01 | 10.0001 | 0.023842 | 38.8331 | 0.98694 | 22.917 | 1.0026 | -0.19569 | 22.917 | 0.044266 | 8.6737e-05 | -4.4846 |  | 0.021833 | -4.5064 | -0.18607
2022-06-10 23:01 | 10.0001 | 0.011921 | 38.5963 | 0.98694 | 22.9588 | 1.0025 | -0.20023 | 22.9588 | -0.1925 | -0.0071331 | -4.597 |  | -0.094939 | -4.5021 | -0.19927
2022-06-11 | 10.0001 | 0.011921 | 38.2561 | 0.98694 | 22.9999 | 0.98591 | -0.20745 | 22.9999 | -0.53276 | -0.15122 | -4.7713 |  | -0.26261 | -4.5087 | -0.19338
2022-06-11 01:00 | 10.0001 | 0.023842 | 38.0918 | 0.98694 | 23.0416 | 1.0025 | -0.21257 | 23.0416 | -0.697 | -0.33861 | -4.8979 |  | -0.34341 | -4.5545 | -0.18906
2022-06-11 02:00 | 10.0001 | 0.0292 | 37.9549 | 0.98694 | 23.0834 | 1.0026 | -0.21403 | 23.0834 | -0.83392 | -0.57993 | -4.9405 |  | -0.41067 | -4.5298 | -0.19785
2022-06-11 03:01 | 10.0001 | 0.020648 | 37.7627 | 0.98694 | 23.1252 | 1.0026 | -0.21855 | 23.1252 | -1.0262 | -1.0805 | -5.054 |  | -0.50493 | -4.5491 | -0.19957
2022-06-11 04:01 | 10.0001 | 0.016859 | 37.5218 | 0.98694 | 23.167 | 1.0025 | -0.22305 | 23.167 | -1.267 | -2.0341 | -5.1675 |  | -0.62268 | -4.5448 | -0.20637
2022-06-11 05:01 | 10.0001 | 0.011921 | 37.3982 | 0.98694 | 23.2087 | 1.0025 | -0.2268 | 23.2087 | -1.3906 | -2.6891 | -5.2637 |  | -0.68289 | -4.5808 | -0.18964
2022-06-11 06:00 | 10.0001 | 0.020648 | 37.3103 | 0.98694 | 23.2498 | 0.98581 | -0.22653 | 23.2498 | -1.4786 | -3.2324 | -5.2668 |  | -0.72567 | -4.5411 | -0.075947
2022-06-11 07:00 | 10.0001 | 0.023842 | 37.4824 | 0.98694 | 23.2916 | 1.0026 | -0.2229 | 23.2916 | -1.3065 | -2.2299 | -5.1918 |  | -0.6419 | -4.5499 | -0.1544
2022-06-11 08:00 | 10.0001 | 0.023842 | 37.699 | 0.98694 | 23.3334 | 1.0026 | -0.21947 | 23.3334 | -1.0898 | -1.2944 | -5.1209 |  | -0.53609 | -4.5848 | -0.1797
2022-06-11 09:01 | 10.0001 | 0.03154 | 37.8909 | 0.98694 | 23.3751 | 1.0026 | -0.21574 | 23.3751 | -0.89795 | -0.72404 | -5.0428 |  | -0.4421 | -4.6007 | -0.20135
2022-06-11 10:01 | 10.0001 | 0.023842 | 38.1784 | 0.98694 | 23.4169 | 1.0026 | -0.20869 | 23.4169 | -0.61038 | -0.2274 | -4.8868 |  | -0.30081 | -4.586 | -0.15321
2022-06-11 11:01 | 10.0001 | 0.03154 | 38.4382 | 0.98694 | 23.4587 | 1.0026 | -0.20387 | 23.4587 | -0.35066 | -0.043118 | -4.7824 |  | -0.17291 | -4.6095 | -0.20783
2022-06-11 12:01 | 10.0001 | 0.016859 | 38.5794 | 0.98694 | 23.5004 | 1.0025 | -0.20122 | 23.5004 | -0.20947 | -0.0091907 | -4.7287 |  | -0.10331 | -4.6254 | -0.24796
2022-06-11 13:00 | 10.0001 | 0.020648 | 38.6682 | 0.98694 | 23.5415 | 0.98585 | -0.1991 | 23.5415 | -0.12066 | -0.0017566 | -4.6872 |  | -0.059512 | -4.6277 | -0.24007
2022-06-11 14:00 | 10.0001 | 0.020648 | 38.7481 | 0.98694 | 23.5833 | 1.0026 | -0.19815 | 23.5833 | -0.040719 | -6.7513e-05 | -4.673 |  | -0.020084 | -4.6529 | -0.19815
2022-06-11 15:00 | 10.0001 | 0.023842 | 38.8087 | 0.98694 | 23.6251 | 1.0025 | -0.19613 | 23.6251 | 0.019852 | 7.8235e-06 | -4.6335 |  | 0.0097917 | -4.6433 | -0.23893
2022-06-11 16:01 | 10.0001 | 0.023842 | 38.8083 | 0.98694 | 23.6668 | 1.0026 | -0.19796 | 23.6668 | 0.019483 | 7.395e-06 | -4.6852 |  | 0.0096096 | -4.6948 | -0.24566
2022-06-11 17:01 | 10.0001 | 0.011921 | 38.8662 | 0.98694 | 23.7086 | 1.0025 | -0.19599 | 23.7086 | 0.07735 | 0.00046279 | -4.6467 |  | 0.038152 | -4.6848 | -0.2919
2022-06-11 18:01 | 10.0001 | 0.023842 | 38.9542 | 0.98694 | 23.7504 | 1.0026 | -0.1936 | 23.7504 | 0.16533 | 0.0045195 | -4.598 |  | 0.081544 | -4.6796 | -0.22977
2022-06-11 19:00 | 10.0001 | 0.011921 | 38.9839 | 0.98694 | 23.7915 | 0.9859 | -0.19152 | 23.7915 | 0.19505 | 0.0074208 | -4.5565 |  | 0.096199 | -4.6527 | -0.18461
2022-06-11 20:00 | 10.0001 | 0.020648 | 39.0784 | 0.98694 | 23.8332 | 1.0025 | -0.18974 | 23.8332 | 0.28956 | 0.024279 | -4.522 |  | 0.1428 | -4.6648 | -0.164
2022-06-11 21:00 | 10.0001 | 0.03154 | 39.1417 | 0.98694 | 23.875 | 1.0025 | -0.18919 | 23.875 | 0.35284 | 0.043928 | -4.517 |  | 0.17399 | -4.691 | -0.19984
2022-06-11 22:01 | 10.0001 | 0.011921 | 38.9816 | 0.98694 | 23.9168 | 1.0026 | -0.1928 | 23.9168 | 0.1928 | 0.0071667 | -4.6112 |  | 0.095088 | -4.7063 | -0.19995
2022-06-11 23:01 | 10.0001 | 0.03154 | 38.6872 | 0.98694 | 23.9586 | 1.0026 | -0.19978 | 23.9586 | -0.1016 | -0.0010487 | -4.7865 |  | -0.050111 | -4.7364 | -0.23436
2022-06-12 00:01 | 10.0001 | 0 | 38.4343 | 0.98694 | 24.0003 | 1.0026 | -0.20404 | 24.0003 | -0.35456 | -0.044573 | -4.8969 |  | -0.17483 | -4.7221 | -0.21332
2022-06-12 01:01 | 10.0001 | 0.033717 | 38.2192 | 0.98694 | 24.0421 | 1.0026 | -0.20849 | 24.0421 | -0.56962 | -0.18482 | -5.0124 |  | -0.28075 | -4.7317 | -0.17707
2022-06-12 02:00 | 10.0001 | NaN | 38.123 | 0.98694 | 24.0832 | 0.9859 | -0.20969 | 24.0832 | -0.66586 | -0.29522 | -5.0499 |  | -0.3281 | -4.7218 | -0.19205
2022-06-12 03:00 | 10.0001 | 0.0292 | 37.8432 | 0.98694 | 24.125 | 1.0026 | -0.21621 | 24.125 | -0.94562 | -0.84558 | -5.2161 |  | -0.46547 | -4.7506 | -0.20162
2022-06-12 04:00 | 10.0001 | 0.011921 | 37.6967 | 0.98694 | 24.1667 | 1.0026 | -0.21898 | 24.1667 | -1.0921 | -1.3024 | -5.2921 |  | -0.5372 | -4.7549 | -0.21015
2022-06-12 05:01 | 10.0001 | 0.039537 | 37.5646 | 0.98694 | 24.2085 | 1.0026 | -0.22049 | 24.2085 | -1.2243 | -1.835 | -5.3377 |  | -0.6018 | -4.7359 | -0.15508
2022-06-12 06:01 | 10.0001 | 0.011921 | 37.4889 | 0.98694 | 24.2503 | 1.0026 | -0.22119 | 24.2503 | -1.2999 | -2.1966 | -5.364 |  | -0.63872 | -4.7253 | -0.18414
2022-06-12 07:01 | 10.0001 | NaN | 37.6659 | 0.98694 | 24.2921 | 1.0026 | -0.2181 | 24.2921 | -1.1229 | -1.416 | -5.2982 |  | -0.55229 | -4.7459 | -0.19591
2022-06-12 08:00 | 10.0001 | 0.020648 | 37.8707 | 0.98694 | 24.3331 | 0.98589 | -0.2137 | 24.3331 | -0.91815 | -0.774 | -5.1999 |  | -0.452 | -4.7479 | -0.16316
2022-06-12 09:00 | 10.0001 | 0.041295 | 38.139 | 0.98694 | 24.3749 | 1.0026 | -0.20801 | 24.3749 | -0.64979 | -0.27436 | -5.0702 |  | -0.3202 | -4.75 | -0.14929
2022-06-12 10:00 | 10.0001 | 0 | 38.3193 | 0.98694 | 24.4167 | 1.0026 | -0.20591 | 24.4167 | -0.46949 | -0.10349 | -5.0276 |  | -0.23146 | -4.7962 | -0.21023
2022-06-12 11:01 | 10.0001 | 0.020648 | 38.4719 | 0.98694 | 24.4585 | 1.0026 | -0.20229 | 24.4585 | -0.31696 | -0.031842 | -4.9476 |  | -0.1563 | -4.7913 | -0.18181
2022-06-12 12:01 | 10.0001 | 0.016859 | 38.5327 | 0.98694 | 24.5002 | 1.0025 | -0.20087 | 24.5002 | -0.25609 | -0.016795 | -4.9213 |  | -0.12629 | -4.795 | -0.16956
2022-06-12 13:01 | 10.0001 | 0.020648 | 38.667 | 0.98694 | 24.542 | 1.0025 | -0.1973 | 24.542 | -0.12179 | -0.0018066 | -4.8422 |  | -0.06007 | -4.7822 | -0.15439
2022-06-12 14:01 | 10.0001 | 0.020648 | 38.8054 | 0.98694 | 24.5838 | 1.0025 | -0.19466 | 24.5838 | 0.016621 | 4.5916e-06 | -4.7855 |  | 0.0081981 | -4.7937 | -0.14068
2022-06-12 15:00 | 10.0001 | 0.033717 | 38.8696 | 0.98694 | 24.6249 | 0.98591 | -0.19429 | 24.6249 | 0.080804 | 0.00052758 | -4.7845 |  | 0.039855 | -4.8243 | -0.18107
2022-06-12 16:00 | 10.0001 | 0.016859 | 38.9151 | 0.98694 | 24.6666 | 1.0026 | -0.19331 | 24.6666 | 0.1263 | 0.0020145 | -4.7683 |  | 0.062292 | -4.8305 | -0.13579
2022-06-12 17:00 | 10.0001 | 0.023842 | 38.9255 | 0.98694 | 24.7084 | 1.0025 | -0.19343 | 24.7084 | 0.13665 | 0.0025517 | -4.7794 |  | 0.067398 | -4.8468 | -0.16201
2022-06-12 18:01 | 10.0001 | 0.026656 | 39.0273 | 0.98694 | 24.7502 | 1.0025 | -0.19098 | 24.7502 | 0.23843 | 0.013555 | -4.7267 |  | 0.11759 | -4.8443 | -0.16476
2022-06-12 19:01 | 10.0001 | 0.011921 | 39.0421 | 0.98694 | 24.792 | 1.0026 | -0.1918 | 24.792 | 0.25323 | 0.016238 | -4.7551 |  | 0.12488 | -4.88 | -0.22717
2022-06-12 20:01 | 10.0001 | 0.020648 | 39.1179 | 0.98694 | 24.8337 | 1.0026 | -0.19021 | 24.8337 | 0.32904 | 0.035623 | -4.7237 |  | 0.16225 | -4.8859 | -0.22098
2022-06-12 21:00 | 10.0001 | 0 | 39.0554 | 0.98694 | 24.8748 | 0.98591 | -0.191 | 24.8748 | 0.26658 | 0.018945 | -4.751 |  | 0.13147 | -4.8825 | -0.19158
2022-06-12 22:00 | 10.0001 | 0.011921 | 38.9237 | 0.98694 | 24.9166 | 1.0026 | -0.19251 | 24.9166 | 0.13484 | 0.0024518 | -4.7966 |  | 0.066507 | -4.8631 | -0.15687
2022-06-12 23:00 | 10.0001 | 0.0292 | 38.9605 | 0.98694 | 24.9584 | 1.0026 | -0.19206 | 24.9584 | 0.17171 | 0.0050625 | -4.7936 |  | 0.084687 | -4.8783 | -0.14192
2022-06-13 00:01 | 10.0001 | 0.020648 | 38.9815 | 0.98694 | 25.0001 | 1.0025 | -0.19146 | 25.0001 | 0.19272 | 0.0071577 | -4.7865 |  | 0.095049 | -4.8815 | -0.15947
2022-06-13 01:01 | 10.0001 | 0.026656 | 38.9213 | 0.98694 | 25.0419 | 1.0026 | -0.19199 | 25.0419 | 0.13252 | 0.0023275 | -4.8078 |  | 0.065363 | -4.8731 | -0.14147
2022-06-13 02:01 | 10.0001 | 0.023842 | 38.8534 | 0.98694 | 25.0837 | 1.0025 | -0.19425 | 25.0837 | 0.064579 | 0.00026932 | -4.8726 |  | 0.031852 | -4.9045 | -0.18255
2022-06-13 03:01 | 10.0001 | 0.011921 | 38.798 | 0.98694 | 25.1254 | 1.0026 | -0.19655 | 25.1254 | 0.0091858 | 7.7509e-07 | -4.9385 |  | 0.0045308 | -4.943 | -0.19228
2022-06-13 04:00 | 10.0001 | 0.020648 | 38.7685 | 0.98694 | 25.1665 | 0.98592 | -0.19402 | 25.1665 | -0.020312 | -8.3808e-06 | -4.8828 |  | -0.010019 | -4.8727 | -0.11786
2022-06-13 05:00 | 10.0001 | 0.016859 | 38.7174 | 0.98694 | 25.2083 | 1.0026 | -0.19711 | 25.2083 | -0.071432 | -0.00036449 | -4.9689 |  | -0.035233 | -4.9336 | -0.1978
2022-06-13 06:00 | 10.0001 | 0.044604 | 38.7305 | 0.98694 | 25.2501 | 1.0026 | -0.19582 | 25.2501 | -0.058368 | -0.00019885 | -4.9446 |  | -0.028789 | -4.9158 | -0.19049
2022-06-13 07:01 | 10.0001 | 0 | 38.8236 | 0.98694 | 25.2919 | 1.0026 | -0.19434 | 25.2919 | 0.03479 | 4.2109e-05 | -4.9152 |  | 0.01716 | -4.9323 | -0.18648
2022-06-13 08:01 | 10.0001 | 0.026656 | 38.9598 | 0.98694 | 25.3336 | 1.0025 | -0.19146 | 25.3336 | 0.17102 | 0.0050021 | -4.8503 |  | 0.084349 | -4.9347 | -0.18669
2022-06-13 09:01 | 10.0001 | 0.011921 | 38.9918 | 0.98694 | 25.3754 | 1.0025 | -0.1907 | 25.3754 | 0.20293 | 0.008357 | -4.8392 |  | 0.10008 | -4.9393 | -0.18917
2022-06-13 10:00 | 10.0001 | 0.020648 | 38.7651 | 0.98694 | 25.4165 | 0.9859 | -0.19426 | 25.4165 | -0.023694 | -1.3303e-05 | -4.9375 |  | -0.011687 | -4.9258 | -0.12962
2022-06-13 11:00 | 10.0001 | 0.016859 | 38.8344 | 0.98694 | 25.4582 | 1.0025 | -0.19414 | 25.4582 | 0.04559 | 9.4757e-05 | -4.9425 |  | 0.022487 | -4.965 | -0.17376
2022-06-13 12:00 | 10.0001 | 0.020648 | 38.7787 | 0.98694 | 25.5 | 1.0026 | -0.19351 | 25.5 | -0.010101 | -1.0305e-06 | -4.9344 |  | -0.004982 | -4.9294 | -0.13451
2022-06-13 13:01 | 10.0001 | 0.011921 | 38.7355 | 0.98694 | 25.5418 | 1.0025 | -0.19466 | 25.5418 | -0.053344 | -0.0001518 | -4.9719 |  | -0.026311 | -4.9456 | -0.16349
2022-06-13 14:01 | 10.0001 | 0.020648 | 39.0605 | 0.98694 | 25.5836 | 1.0025 | -0.18923 | 25.5836 | 0.27169 | 0.020055 | -4.8412 |  | 0.13398 | -4.9752 | -0.18155
2022-06-13 15:01 | 10.0001 | 0.023842 | 39.3564 | 0.98694 | 25.6253 | 1.0025 | -0.18343 | 25.6253 | 0.56759 | 0.18285 | -4.7004 |  | 0.27975 | -4.9801 | -0.15573
2022-06-13 16:01 | 10.0001 | 0.016859 | 39.6697 | 0.98694 | 25.6671 | 1.0027 | -0.17886 | 25.6671 | 0.88083 | 0.68341 | -4.5909 |  | 0.4337 | -5.0246 | -0.194
2022-06-13 18:01 | 10.0001 | 0.011921 | 39.7965 | 0.98694 | 25.7504 | 1.9976 | -0.18124 | 25.7504 | 1.0077 | 1.0232 | -4.6669 |  | 0.49589 | -5.1628 | -0.3033
2022-06-13 19:02 | 10.0001 | 0.020648 | 40.0934 | 0.98694 | 25.7926 | 1.0139 | -0.17595 | 25.7926 | 1.3046 | 2.2202 | -4.5383 |  | 0.64098 | -5.1792 | -0.3213
2022-06-13 20:01 | 10.0001 | 0.016859 | 40.1366 | 0.98694 | 25.8337 | 0.98592 | -0.17414 | 25.8337 | 1.3477 | 2.448 | -4.4987 |  | 0.66202 | -5.1608 | -0.26954
2022-06-13 21:00 | 10.0001 | 0.0292 | 39.7435 | 0.98694 | 25.8748 | 0.98597 | -0.1808 | 25.8748 | 0.95469 | 0.87015 | -4.6781 |  | 0.46992 | -5.148 | -0.25174
2022-06-13 22:00 | 10.0001 | 0.0292 | 39.3587 | 0.98694 | 25.9165 | 1.0027 | -0.18821 | 25.9165 | 0.56984 | 0.18503 | -4.8777 |  | 0.28086 | -5.1585 | -0.26499
2022-06-13 23:00 | 10.0001 | 0.016859 | 39.2124 | 0.98694 | 25.9583 | 1.0026 | -0.19068 | 25.9583 | 0.42354 | 0.075975 | -4.9496 |  | 0.20882 | -5.1584 | -0.28348
2022-06-14 | 10.0001 | 0.026656 | 39.0464 | 0.98694 | 26.0001 | 1.0026 | -0.19337 | 26.0001 | 0.25758 | 0.017089 | -5.0276 |  | 0.12703 | -5.1547 | -0.26532
2022-06-14 01:01 | 10.0001 | 0.020648 | 38.9346 | 0.98694 | 26.0419 | 1.0027 | -0.19676 | 26.0419 | 0.14579 | 0.0030988 | -5.1239 |  | 0.071906 | -5.1958 | -0.30167
2022-06-14 09:00 | 10.0001 | 0.037697 | 38.5145 | 0.98694 | 26.3749 | 7.992 | -0.20299 | 26.3749 | -0.27434 | -0.020647 | -5.3539 |  | -0.13529 | -5.2186 | -0.25918
2022-06-14 10:00 | 10.0001 | 0.020648 | 37.8124 | 0.98694 | 26.4166 | 1.0026 | -0.21218 | 26.4166 | -0.97641 | -0.93089 | -5.6052 |  | -0.48056 | -5.1246 | -0.16516
2022-06-14 11:00 | 10.0001 | 0.023842 | 37.8295 | 0.98694 | 26.4584 | 1.0026 | -0.21204 | 26.4584 | -0.95929 | -0.88278 | -5.6102 |  | -0.47217 | -5.1381 | -0.14634
2022-06-14 12:01 | 10.0001 | 0 | 38.0929 | 0.98694 | 26.5002 | 1.0026 | -0.20769 | 26.5002 | -0.69597 | -0.33712 | -5.5039 |  | -0.3429 | -5.161 | -0.21611
2022-06-14 13:01 | 10.0001 | 0.026656 | 38.2224 | 0.98694 | 26.542 | 1.0026 | -0.20477 | 26.542 | -0.56641 | -0.18171 | -5.435 |  | -0.27917 | -5.1558 | -0.16657
2022-06-14 14:01 | 10.0001 | 0.026656 | 38.1794 | 0.98694 | 26.5837 | 1.0027 | -0.20502 | 26.5837 | -0.60943 | -0.22634 | -5.4502 |  | -0.30034 | -5.1498 | -0.17549
2022-06-14 15:00 | 10.0001 | 0.020648 | 38.2636 | 0.98694 | 26.6248 | 0.98594 | -0.20341 | 26.6248 | -0.5252 | -0.14487 | -5.4157 |  | -0.25889 | -5.1568 | -0.16841
2022-06-14 16:00 | 10.0001 | 0.023842 | 38.461 | 0.98694 | 26.6666 | 1.0026 | -0.19861 | 26.6666 | -0.32786 | -0.035244 | -5.2961 |  | -0.16168 | -5.1345 | -0.14989
2022-06-14 17:00 | 10.0001 | 0.011921 | 38.334 | 0.98694 | 26.7084 | 1.0026 | -0.20038 | 26.7084 | -0.4548 | -0.09407 | -5.3519 |  | -0.22422 | -5.1277 | -0.14133
2022-06-14 18:01 | 10.0001 | NaN | 38.4876 | 0.98694 | 26.7502 | 1.0026 | -0.20083 | 26.7502 | -0.30122 | -0.027329 | -5.3721 |  | -0.14854 | -5.2236 | -0.22331
2022-06-14 19:01 | 10.0001 | 0.033717 | 39.1421 | 0.98694 | 26.7919 | 1.0026 | -0.18849 | 26.7919 | 0.35328 | 0.044091 | -5.0499 |  | 0.1742 | -5.2241 | -0.19429
2022-06-14 20:01 | 10.0001 | 0.020648 | 39.2981 | 0.98694 | 26.8337 | 1.0026 | -0.18336 | 26.8337 | 0.50926 | 0.13208 | -4.9202 |  | 0.25104 | -5.1713 | -0.18133
2022-06-14 21:00 | 10.0001 | NaN | 37.8539 | 0.98694 | 26.8748 | 0.98595 | -0.20985 | 26.8748 | -0.9349 | -0.81715 | -5.6396 |  | -0.46022 | -5.1794 | -0.1754
2022-06-14 22:00 | 10.0001 | 0.042982 | 38.7419 | 0.98694 | 26.9166 | 1.0027 | -0.1939 | 26.9166 | -0.046893 | -0.00010311 | -5.2191 |  | -0.023129 | -5.196 | -0.16566
2022-06-14 23:00 | 10.0001 | 0.023842 | 39.2751 | 0.98694 | 26.9583 | 1.0026 | -0.18597 | 26.9583 | 0.48629 | 0.115 | -5.0135 |  | 0.23973 | -5.2532 | -0.20485
2022-06-15 00:01 | 10.0001 | 0.016859 | 39.1845 | 0.98694 | 27.0001 | 1.0026 | -0.18895 | 27.0001 | 0.39569 | 0.061952 | -5.1016 |  | 0.1951 | -5.2967 | -0.2041
2022-06-15 01:01 | 10.0001 | 0.0292 | 38.9629 | 0.98694 | 27.0419 | 1.0026 | -0.19252 | 27.0419 | 0.1741 | 0.0052767 | -5.206 |  | 0.085864 | -5.2918 | -0.099894
2022-06-15 02:01 | 10.0001 | 0.023842 | 38.8397 | 0.98694 | 27.0837 | 1.0026 | -0.19353 | 27.0837 | 0.050839 | 0.0001314 | -5.2414 |  | 0.025076 | -5.2665 | -0.067599
2022-06-15 03:01 | 10.0001 | 0.037697 | 38.7753 | 0.98694 | 27.1254 | 1.0027 | -0.19562 | 27.1254 | -0.013563 | -2.4949e-06 | -5.3063 |  | -0.0066897 | -5.2996 | -0.10748
2022-06-15 04:00 | 10.0001 | NaN | 38.8426 | 0.98694 | 27.1665 | 0.98599 | -0.1938 | 27.1665 | 0.053764 | 0.00015541 | -5.2647 |  | 0.026518 | -5.2913 | -0.1109
2022-06-15 05:00 | 10.0001 | 0.023842 | 38.7667 | 0.98694 | 27.2083 | 1.0026 | -0.19473 | 27.2083 | -0.022127 | -1.0834e-05 | -5.2982 |  | -0.010914 | -5.2873 | -0.099645
2022-06-15 06:00 | 10.0001 | 0.016859 | 38.7443 | 0.98694 | 27.2501 | 1.0026 | -0.19588 | 27.2501 | -0.044494 | -8.8086e-05 | -5.3377 |  | -0.021946 | -5.3157 | -0.12177
2022-06-15 07:01 | 10.0001 | 0.026656 | 39.0253 | 0.98694 | 27.2919 | 1.0026 | -0.19057 | 27.2919 | 0.23648 | 0.013224 | -5.2009 |  | 0.11662 | -5.3175 | -0.12608
2022-06-15 08:01 | 10.0001 | 0.03154 | 39.2068 | 0.98694 | 27.3336 | 1.0026 | -0.18812 | 27.3336 | 0.41798 | 0.073022 | -5.1421 |  | 0.20608 | -5.3482 | -0.11799
2022-06-15 09:01 | 10.0001 | 0.020648 | 39.3668 | 0.98694 | 27.3754 | 1.0026 | -0.18362 | 27.3754 | 0.57794 | 0.19304 | -5.0266 |  | 0.28484 | -5.3115 | -0.092813
2022-06-15 10:00 | 10.0001 | 0 | 39.2059 | 0.98694 | 27.4165 | 0.98594 | -0.18818 | 27.4165 | 0.41707 | 0.07255 | -5.1594 |  | 0.20564 | -5.365 | -0.24041
2022-06-15 11:00 | 10.0001 | 0.016859 | 39.2359 | 0.98694 | 27.4583 | 1.0026 | -0.1883 | 27.4583 | 0.4471 | 0.089372 | -5.1705 |  | 0.22042 | -5.3909 | -0.2529
2022-06-15 12:00 | 10.0001 | 0.0292 | 39.3734 | 0.98694 | 27.5 | 1.0026 | -0.18419 | 27.5 | 0.58462 | 0.19981 | -5.0651 |  | 0.28813 | -5.3533 | -0.19234
2022-06-15 13:01 | 10.0001 | NaN | 39.6279 | 0.98694 | 27.5418 | 1.0027 | -0.17964 | 27.5418 | 0.83908 | 0.59076 | -4.9476 |  | 0.41321 | -5.3608 | -0.20505
2022-06-15 14:01 | 10.0001 | 0.011921 | 39.6022 | 0.98694 | 27.5836 | 1.0026 | -0.1822 | 27.5836 | 0.81334 | 0.53805 | -5.0256 |  | 0.40057 | -5.4262 | -0.27641
2022-06-15 15:01 | 10.0001 | 0.026656 | 40.1746 | 0.98694 | 27.6254 | 1.0026 | -0.17117 | 27.6254 | 1.3858 | 2.6612 | -4.7287 |  | 0.68054 | -5.4093 | -0.25233
2022-06-15 16:00 | 10.0001 | 0.026656 | 40.0248 | 0.98694 | 27.6665 | 0.98597 | -0.17506 | 27.6665 | 1.236 | 1.8883 | -4.8432 |  | 0.60754 | -5.4508 | -0.31635
2022-06-15 17:00 | 10.0001 | 0.0292 | 39.8803 | 0.98694 | 27.7082 | 1.0026 | -0.17812 | 27.7082 | 1.0915 | 1.3005 | -4.9354 |  | 0.53693 | -5.4724 | -0.34476
2022-06-15 18:00 | 10.0001 | 0.011921 | 39.7974 | 0.98694 | 27.75 | 1.0026 | -0.17997 | 27.75 | 1.0086 | 1.026 | -4.9942 |  | 0.49632 | -5.4905 | -0.26697
2022-06-15 19:01 | 10.0001 | 0.035763 | 39.4362 | 0.98694 | 27.7918 | 1.0026 | -0.18134 | 27.7918 | 0.64738 | 0.27131 | -5.0398 |  | 0.31901 | -5.3588 | -0.13469
2022-06-15 20:01 | 10.0001 | 0.026656 | 38.6481 | 0.98694 | 27.8336 | 1.0026 | -0.19792 | 27.8336 | -0.14074 | -0.0027876 | -5.5089 |  | -0.069414 | -5.4395 | -0.26827
2022-06-15 21:01 | 10.0001 | 0.016859 | 38.9144 | 0.98694 | 27.8753 | 1.0026 | -0.19265 | 27.8753 | 0.1256 | 0.0019813 | -5.3701 |  | 0.061947 | -5.4321 | -0.25251
2022-06-15 22:01 | 10.0001 | 0.0292 | 39.1632 | 0.98694 | 27.9171 | 1.0026 | -0.1888 | 27.9171 | 0.37436 | 0.052464 | -5.2708 |  | 0.18459 | -5.4554 | -0.25925
2022-06-15 23:00 | 10.0001 | 0.011921 | 39.1818 | 0.98694 | 27.9582 | 0.98594 | -0.19005 | 27.9582 | 0.39298 | 0.060689 | -5.3134 |  | 0.19376 | -5.5071 | -0.25399
2022-06-16 | 10.0001 | 0.011921 | 39.3152 | 0.98694 | 28 | 1.0026 | -0.18687 | 28 | 0.52636 | 0.14583 | -5.2323 |  | 0.25946 | -5.4918 | -0.1951
2022-06-16 01:00 | 10.0001 | 0.023842 | 39.3224 | 0.98694 | 28.0417 | 1.0026 | -0.18572 | 28.0417 | 0.53357 | 0.15191 | -5.208 |  | 0.26301 | -5.471 | -0.1792
2022-06-16 02:01 | 10.0001 | 0 | 39.1711 | 0.98694 | 28.0835 | 1.0026 | -0.18772 | 28.0835 | 0.38224 | 0.055847 | -5.2718 |  | 0.18847 | -5.4603 | -0.19382
2022-06-16 03:01 | 10.0001 | 0.03154 | 38.9179 | 0.98694 | 28.1253 | 1.0026 | -0.1914 | 28.1253 | 0.12912 | 0.0021526 | -5.3833 |  | 0.063684 | -5.447 | -0.1474
2022-06-16 04:01 | 10.0001 | 0.011921 | 38.8084 | 0.98694 | 28.1671 | 1.0026 | -0.19339 | 28.1671 | 0.019526 | 7.4441e-06 | -5.4471 |  | 0.0096308 | -5.4568 | -0.16541
2022-06-16 05:00 | 10.0001 | 0.0292 | 38.7844 | 0.98694 | 28.2081 | 0.9859 | -0.19407 | 28.2081 | -0.0044229 | -8.6521e-08 | -5.4745 |  | -0.0021815 | -5.4723 | -0.18507
2022-06-16 06:00 | 10.0001 | NaN | 38.8314 | 0.98694 | 28.2499 | 1.0027 | -0.19156 | 28.2499 | 0.042575 | 7.7174e-05 | -5.4117 |  | 0.021 | -5.4327 | -0.11693
2022-06-16 07:00 | 10.0001 | 0.026656 | 38.9518 | 0.98694 | 28.2917 | 1.0026 | -0.19089 | 28.2917 | 0.16298 | 0.0043293 | -5.4005 |  | 0.080384 | -5.4809 | -0.16339
2022-06-16 08:01 | 10.0001 | 0.023842 | 39.0688 | 0.98694 | 28.3335 | 1.0026 | -0.18857 | 28.3335 | 0.27993 | 0.021936 | -5.3428 |  | 0.13805 | -5.4808 | -0.13261
2022-06-16 09:01 | 10.0001 | 0.016859 | 38.6205 | 0.98694 | 28.3752 | 1.0026 | -0.19457 | 28.3752 | -0.16836 | -0.0047719 | -5.5211 |  | -0.083035 | -5.4381 | -0.1266
2022-06-16 10:01 | 10.0001 | 0.011921 | 39.1177 | 0.98694 | 28.417 | 1.0026 | -0.18716 | 28.417 | 0.32887 | 0.035568 | -5.3184 |  | 0.16217 | -5.4806 | -0.11555
2022-06-16 11:00 | 10.0001 | 0.016859 | 39.4571 | 0.98694 | 28.4581 | 0.98595 | -0.17866 | 28.4581 | 0.66824 | 0.2984 | -5.0844 |  | 0.32927 | -5.4137 | -0.022723
2022-06-16 12:00 | 10.0001 | 0 | 39.6214 | 0.98694 | 28.4999 | 1.0026 | -0.17559 | 28.4999 | 0.83255 | 0.57708 | -5.0043 |  | 0.41 | -5.4143 | -0.061085
2022-06-16 13:00 | 10.0001 | 0.0292 | 39.741 | 0.98694 | 28.5417 | 1.0027 | -0.17647 | 28.5417 | 0.95213 | 0.86315 | -5.0368 |  | 0.46866 | -5.5054 | -0.14464
2022-06-16 14:00 | 10.0001 | 0.026656 | 40.0582 | 0.98694 | 28.5834 | 1.0027 | -0.16912 | 28.5834 | 1.2694 | 2.0453 | -4.8341 |  | 0.62381 | -5.4579 | -0.031742
2022-06-16 15:01 | 10.0001 | 0.049151 | 40.4815 | 0.98694 | 28.6252 | 1.0026 | -0.1624 | 28.6252 | 1.6926 | 4.8494 | -4.6487 |  | 0.82945 | -5.4781 | -0.06887
2022-06-16 16:01 | 10.0001 | 0.020648 | 40.8226 | 0.98694 | 28.667 | 1.0026 | -0.15485 | 28.667 | 2.0337 | 8.4117 | -4.439 |  | 0.99371 | -5.4327 | 0.018099
2022-06-16 17:01 | 10.0001 | 0.03154 | 41.2386 | 0.98694 | 28.7088 | 1.0026 | -0.14989 | 28.7088 | 2.4498 | 14.7028 | -4.3032 |  | 1.1919 | -5.4951 | -0.02271
2022-06-16 18:00 | 10.0001 | 0.020648 | 41.5511 | 0.98694 | 28.7498 | 0.9859 | -0.14432 | 28.7498 | 2.7622 | 21.0756 | -4.1492 |  | 1.3389 | -5.488 | 0.0024813
2022-06-16 19:00 | 10.0001 | 0.023842 | 41.4467 | 0.98694 | 28.7916 | 1.0026 | -0.14763 | 28.7916 | 2.6579 | 18.7764 | -4.2505 |  | 1.29 | -5.5405 | -0.1817
2022-06-16 20:00 | 10.0001 | 0.023842 | 41.0072 | 0.98694 | 28.8334 | 1.0027 | -0.15402 | 28.8334 | 2.2184 | 10.9173 | -4.441 |  | 1.082 | -5.523 | -0.083474
2022-06-16 21:01 | 10.0001 | 0.033717 | 40.5218 | 0.98694 | 28.8752 | 1.0026 | -0.16184 | 28.8752 | 1.7329 | 5.2042 | -4.673 |  | 0.84894 | -5.5219 | -0.089902
2022-06-16 22:01 | 10.0001 | 0.023842 | 40.1113 | 0.98694 | 28.9169 | 1.0027 | -0.17022 | 28.9169 | 1.3225 | 2.313 | -4.9223 |  | 0.64972 | -5.572 | -0.1166
2022-06-16 23:01 | 10.0001 | 0.020648 | 40.0082 | 0.98694 | 28.9587 | 1.0026 | -0.17312 | 28.9587 | 1.2194 | 1.8133 | -5.0135 |  | 0.59944 | -5.6129 | -0.1057
2022-06-17 | 10.0001 | 0.023842 | 39.8291 | 0.98694 | 28.9998 | 0.9859 | -0.17749 | 28.9998 | 1.0403 | 1.1257 | -5.1472 |  | 0.51183 | -5.659 | -0.16729
2022-06-17 01:00 | 10.0001 | 0.026656 | 39.7419 | 0.98694 | 29.0416 | 1.0026 | -0.17978 | 29.0416 | 0.95309 | 0.86577 | -5.2212 |  | 0.46913 | -5.6903 | -0.21933
2022-06-17 02:00 | 10.0001 | 0.020648 | 39.5986 | 0.98694 | 29.0833 | 1.0026 | -0.18193 | 29.0833 | 0.80979 | 0.53103 | -5.2911 |  | 0.39883 | -5.6899 | -0.22964
2022-06-17 03:01 | 10.0001 | 0.020648 | 39.3224 | 0.98694 | 29.1251 | 1.0026 | -0.18358 | 29.1251 | 0.53357 | 0.15191 | -5.3468 |  | 0.26301 | -5.6098 | -0.16288
2022-06-17 04:01 | 10.0001 | 0.020648 | 39.1135 | 0.98694 | 29.1669 | 1.0026 | -0.18759 | 29.1669 | 0.32466 | 0.034221 | -5.4714 |  | 0.1601 | -5.6315 | -0.17481
2022-06-17 05:01 | 10.0001 | 0.035763 | 39.042 | 0.98694 | 29.2087 | 1.0026 | -0.18923 | 29.2087 | 0.2532 | 0.016232 | -5.5272 |  | 0.12487 | -5.652 | -0.17964
2022-06-17 06:01 | 10.0001 | 0.016859 | 38.9959 | 0.98694 | 29.2505 | 1.0027 | -0.18993 | 29.2505 | 0.20711 | 0.0088839 | -5.5555 |  | 0.10214 | -5.6577 | -0.2249
2022-06-17 07:00 | 10.0001 | 0.023842 | 38.9406 | 0.98694 | 29.2915 | 0.98594 | -0.19056 | 29.2915 | 0.15179 | 0.0034975 | -5.5819 |  | 0.074866 | -5.6567 | -0.17588
2022-06-17 08:00 | 10.0001 | 0.016859 | 38.8534 | 0.98694 | 29.3333 | 1.0026 | -0.19205 | 29.3333 | 0.064561 | 0.0002691 | -5.6336 |  | 0.031844 | -5.6654 | -0.18462
2022-06-17 09:00 | 10.0001 | 0.020648 | 38.9131 | 0.98694 | 29.3751 | 1.0026 | -0.19109 | 29.3751 | 0.12423 | 0.0019175 | -5.6133 |  | 0.061275 | -5.6746 | -0.23655
2022-06-17 10:01 | 10.0001 | 0.016859 | 39.5712 | 0.98694 | 29.4169 | 1.0026 | -0.17976 | 29.4169 | 0.7824 | 0.47894 | -5.288 |  | 0.38537 | -5.6734 | -0.19284
2022-06-17 11:01 | 10.0001 | 0.011921 | 39.8996 | 0.98694 | 29.4586 | 1.0027 | -0.17579 | 29.4586 | 1.1107 | 1.3703 | -5.1786 |  | 0.54632 | -5.7249 | -0.31112
2022-06-17 12:01 | 10.0001 | 0.03154 | 40.0581 | 0.98694 | 29.5004 | 1.0026 | -0.1728 | 29.5004 | 1.2693 | 2.0449 | -5.0976 |  | 0.62377 | -5.7213 | -0.30682
2022-06-17 13:00 | 10.0001 | 0.016859 | 40.1928 | 0.98694 | 29.5415 | 0.98592 | -0.17046 | 29.5415 | 1.404 | 2.7676 | -5.0357 |  | 0.68942 | -5.7252 | -0.21978
2022-06-17 14:00 | 10.0001 | 0.020648 | 40.4405 | 0.98694 | 29.5833 | 1.0026 | -0.16796 | 29.5833 | 1.6516 | 4.5055 | -4.9689 |  | 0.80962 | -5.7785 | -0.32062
2022-06-17 15:00 | 10.0001 | 0.026656 | 40.7212 | 0.98694 | 29.625 | 1.0027 | -0.16311 | 29.625 | 1.9324 | 7.2157 | -4.8321 |  | 0.94506 | -5.7771 | -0.29905
2022-06-17 16:01 | 10.0001 | 0.026656 | 40.8751 | 0.98694 | 29.6668 | 1.0026 | -0.15994 | 29.6668 | 2.0863 | 9.0806 | -4.745 |  | 1.0189 | -5.7638 | -0.33121
2022-06-17 17:01 | 10.0001 | 0 | 41.1112 | 0.98694 | 29.7086 | 1.0026 | -0.15695 | 29.7086 | 2.3223 | 12.5251 | -4.6629 |  | 1.1315 | -5.7943 | -0.29931
2022-06-17 18:01 | 10.0001 | 0.023842 | 41.248 | 0.98694 | 29.7504 | 1.0027 | -0.15336 | 29.7504 | 2.4592 | 14.8722 | -4.5626 |  | 1.1963 | -5.7589 | -0.27071
2022-06-17 19:00 | 10.0001 | 0.023842 | 41.4692 | 0.98694 | 29.7915 | 0.98594 | -0.15159 | 29.7915 | 2.6804 | 19.2575 | -4.516 |  | 1.3005 | -5.8165 | -0.27607
2022-06-17 20:00 | 10.0001 | 0.033717 | 41.5866 | 0.98694 | 29.8332 | 1.0026 | -0.14618 | 29.8332 | 2.7978 | 21.9004 | -4.3609 |  | 1.3555 | -5.7164 | -0.1935
2022-06-17 21:00 | 10.0001 | 0.03154 | 41.6667 | 0.98694 | 29.875 | 1.0026 | -0.14753 | 29.875 | 2.8779 | 23.8355 | -4.4075 |  | 1.3928 | -5.8004 | -0.27848
2022-06-17 22:01 | 10.0001 | 0.023842 | 41.6691 | 0.98694 | 29.9168 | 1.0026 | -0.14736 | 29.9168 | 2.8803 | 23.8953 | -4.4086 |  | 1.394 | -5.8025 | -0.23057
2022-06-17 23:01 | 10.0001 | 0.03154 | 41.5567 | 0.98694 | 29.9586 | 1.0026 | -0.15013 | 29.9586 | 2.7679 | 21.206 | -4.4977 |  | 1.3415 | -5.8393 | -0.2264
2022-06-18 00:01 | 10.0001 | 0.023842 | 41.4993 | 0.98694 | 30.0003 | 1.0026 | -0.15036 | 30.0003 | 2.7105 | 19.9136 | -4.5109 |  | 1.3147 | -5.8256 | -0.16643
2022-06-18 01:01 | 10.0001 | 0.011921 | 41.4649 | 0.98694 | 30.0421 | 1.0027 | -0.151 | 30.0421 | 2.6761 | 19.1653 | -4.5362 |  | 1.2985 | -5.8348 | -0.14438
2022-06-18 02:00 | 10.0001 | 0.020648 | 41.4899 | 0.98694 | 30.0832 | 0.986 | -0.15076 | 30.0832 | 2.7011 | 19.7074 | -4.5352 |  | 1.3103 | -5.8455 | -0.15559
2022-06-18 03:00 | 10.0001 | 0.026656 | 41.4008 | 0.98694 | 30.125 | 1.0026 | -0.15166 | 30.125 | 2.6119 | 17.8192 | -4.5687 |  | 1.2684 | -5.837 | -0.22725
2022-06-18 04:00 | 10.0001 | 0.033717 | 41.1633 | 0.98694 | 30.1667 | 1.0027 | -0.15662 | 30.1667 | 2.3745 | 13.3881 | -4.7247 |  | 1.1562 | -5.8809 | -0.24942
2022-06-18 05:01 | 10.0001 | 0 | 40.8758 | 0.98694 | 30.2085 | 1.0026 | -0.15865 | 30.2085 | 2.0869 | 9.0893 | -4.7926 |  | 1.0192 | -5.8118 | -0.15976
2022-06-18 06:01 | 10.0001 | 0.033717 | 40.2785 | 0.98694 | 30.2503 | 1.0025 | -0.17143 | 30.2503 | 1.4896 | 3.3056 | -5.1857 |  | 0.73105 | -5.9168 | -0.25912
2022-06-18 07:01 | 10.0001 | 0.026656 | 40.0539 | 0.98694 | 30.2921 | 1.0027 | -0.17323 | 30.2921 | 1.2651 | 2.0246 | -5.2475 |  | 0.62171 | -5.8692 | -0.21236
2022-06-18 08:00 | 10.0001 | 0 | 39.9516 | 0.98694 | 30.3331 | 0.98597 | -0.1752 | 30.3331 | 1.1627 | 1.572 | -5.3144 |  | 0.57175 | -5.8861 | -0.22078
2022-06-18 09:00 | 10.0001 | 0.023842 | 40.0576 | 0.98694 | 30.3749 | 1.0026 | -0.17146 | 30.3749 | 1.2688 | 2.0426 | -5.208 |  | 0.62354 | -5.8315 | -0.157
2022-06-18 10:00 | 10.0001 | 0.026656 | 40.2418 | 0.98694 | 30.4167 | 1.0026 | -0.16969 | 30.4167 | 1.4529 | 3.0672 | -5.1614 |  | 0.71322 | -5.8746 | -0.20122
2022-06-18 11:01 | 10.0001 | 0 | 40.4218 | 0.98694 | 30.4585 | 1.0026 | -0.16806 | 30.4585 | 1.633 | 4.3548 | -5.1188 |  | 0.8006 | -5.9194 | -0.19452
2022-06-18 12:01 | 10.0001 | 0.020648 | 40.392 | 0.98694 | 30.5003 | 1.0026 | -0.16647 | 30.5003 | 1.6031 | 4.1202 | -5.0773 |  | 0.78613 | -5.8634 | -0.14212
2022-06-18 13:01 | 10.0001 | 0 | 40.2005 | 0.98694 | 30.542 | 1.0026 | -0.17005 | 30.542 | 1.4117 | 2.8131 | -5.1938 |  | 0.69314 | -5.8869 | -0.1617
2022-06-18 14:00 | 10.0001 | 0.016859 | 39.9489 | 0.98694 | 30.5831 | 0.98597 | -0.17546 | 30.5831 | 1.1601 | 1.5614 | -5.3661 |  | 0.57047 | -5.9365 | -0.15806
2022-06-18 15:00 | 10.0001 | 0.026656 | 39.3573 | 0.98694 | 30.6249 | 1.0026 | -0.18074 | 30.6249 | 0.56846 | 0.1837 | -5.5353 |  | 0.28018 | -5.8155 | -0.03831
2022-06-18 16:00 | 10.0001 | 0.016859 | 38.8451 | 0.98694 | 30.6667 | 1.0026 | -0.19087 | 30.6667 | 0.056293 | 0.00017839 | -5.8534 |  | 0.027766 | -5.8812 | -0.11738
2022-06-18 17:00 | 10.0001 | 0.020648 | 38.1716 | 0.98694 | 30.7084 | 1.0026 | -0.19843 | 30.7084 | -0.61718 | -0.23509 | -6.0936 |  | -0.30415 | -5.7894 | 0.0049442
2022-06-18 18:01 | 10.0001 | 0.016859 | 37.646 | 0.98694 | 30.7502 | 1.0026 | -0.20805 | 30.7502 | -1.1429 | -1.4927 | -6.3975 |  | -0.56203 | -5.8355 | -0.076602
2022-06-18 19:01 | 10.0001 | 0.020648 | 37.5238 | 0.98694 | 30.792 | 1.0026 | -0.21014 | 30.792 | -1.265 | -2.0242 | -6.4705 |  | -0.62168 | -5.8488 | -0.03228
2022-06-18 20:01 | 10.0001 | 0.011921 | 37.5771 | 0.98694 | 30.8338 | 1.0026 | -0.21024 | 30.8338 | -1.2117 | -1.779 | -6.4826 |  | -0.59566 | -5.887 | -0.17045
2022-06-18 21:00 | 10.0001 | 0.033717 | 36.9407 | 0.98694 | 30.8749 | 0.986 | -0.21922 | 30.8749 | -1.8482 | -6.3127 | -6.7684 |  | -0.90452 | -5.8638 | -0.06347
2022-06-18 22:00 | 10.0001 | 0.020648 | 37.0891 | 0.98694 | 30.9166 | 1.0027 | -0.21768 | 30.9166 | -1.6998 | -4.9109 | -6.7299 |  | -0.83289 | -5.897 | -0.094469
2022-06-18 23:00 | 10.0001 | 0.011921 | 37.2738 | 0.98694 | 30.9584 | 1.0027 | -0.21493 | 30.9584 | -1.515 | -3.4775 | -6.6539 |  | -0.74338 | -5.9105 | -0.071246
2022-06-19 00:01 | 10.0001 | 0.016859 | 37.0177 | 0.98694 | 31.0002 | 1.0027 | -0.21621 | 31.0002 | -1.7711 | -5.556 | -6.7025 |  | -0.86738 | -5.8351 | -0.0095714
2022-06-19 01:01 | 10.0001 | 0.016859 | 36.6117 | 0.98694 | 31.042 | 1.0026 | -0.2228 | 31.042 | -2.1771 | -10.3187 | -6.9163 |  | -1.0623 | -5.854 | -0.019256
2022-06-19 02:01 | 10.0001 | 0.023842 | 36.2745 | 0.98694 | 31.0837 | 1.0026 | -0.22818 | 31.0837 | -2.5143 | -15.8943 | -7.0926 |  | -1.2224 | -5.8702 | -0.024745
2022-06-19 03:00 | 10.0001 | 0.0292 | 35.8756 | 0.98694 | 31.1248 | 0.98588 | -0.23475 | 31.1248 | -2.9132 | -24.7247 | -7.3064 |  | -1.4093 | -5.8971 | -0.060087
2022-06-19 04:00 | 10.0001 | 0.026656 | 35.7227 | 0.98694 | 31.1666 | 1.0026 | -0.23723 | 31.1666 | -3.0662 | -28.8258 | -7.3935 |  | -1.4801 | -5.9134 | -0.032505
2022-06-19 05:00 | 10.0001 | 0.011921 | 35.873 | 0.98694 | 31.2084 | 1.0026 | -0.23457 | 31.2084 | -2.9159 | -24.7916 | -7.3206 |  | -1.4105 | -5.9101 | -0.09832
2022-06-19 06:01 | 10.0001 | 0.04617 | 36.1673 | 0.98694 | 31.2501 | 1.0026 | -0.2281 | 31.2501 | -2.6215 | -18.0164 | -7.1281 |  | -1.2729 | -5.8552 | 0.061604
2022-06-19 07:01 | 10.0001 | 0.020648 | 36.1644 | 0.98694 | 31.2919 | 1.0027 | -0.22802 | 31.2919 | -2.6244 | -18.0755 | -7.1352 |  | -1.2742 | -5.8609 | 0.0083113
2022-06-19 08:01 | 10.0001 | 0.016859 | 35.6451 | 0.98694 | 31.3337 | 1.0026 | -0.23567 | 31.3337 | -3.1437 | -31.0684 | -7.3844 |  | -1.5159 | -5.8686 | 0.01757
2022-06-19 09:00 | 10.0001 | 0.020648 | 35.6981 | 0.98694 | 31.3748 | 0.98596 | -0.23407 | 31.3748 | -3.0908 | -29.5258 | -7.3439 |  | -1.4915 | -5.8524 | -0.020869
2022-06-19 10:00 | 10.0001 | 0.023842 | 35.5588 | 0.98694 | 31.4166 | 1.0026 | -0.23702 | 31.4166 | -3.23 | -33.6986 | -7.4462 |  | -1.5555 | -5.8907 | -0.016127
2022-06-19 11:00 | 10.0001 | 0 | 35.7588 | 0.98694 | 31.4583 | 1.0026 | -0.23525 | 31.4583 | -3.03 | -27.8175 | -7.4006 |  | -1.4634 | -5.9372 | -0.017795
2022-06-19 12:00 | 10.0001 | 0.020648 | 36.314 | 0.98694 | 31.5001 | 1.0026 | -0.22523 | 31.5001 | -2.4748 | -15.1572 | -7.0946 |  | -1.2037 | -5.8909 | -0.027496
2022-06-19 13:01 | 10.0001 | 0.03154 | 36.2587 | 0.98694 | 31.5419 | 1.0026 | -0.22512 | 31.5419 | -2.5301 | -16.1964 | -7.1007 |  | -1.2298 | -5.8709 | 0.016081
2022-06-19 14:01 | 10.0001 | 0.044604 | 36.4341 | 0.98694 | 31.5837 | 1.0026 | -0.22376 | 31.5837 | -2.3547 | -13.0558 | -7.0673 |  | -1.1468 | -5.9204 | 0.016076
2022-06-19 15:01 | 10.0001 | 0 | 36.5297 | 0.98694 | 31.6254 | 1.0026 | -0.22177 | 31.6254 | -2.2591 | -11.53 | -7.0136 |  | -1.1014 | -5.9122 | -0.096664
2022-06-19 16:00 | 10.0001 | 0.035763 | 36.5704 | 0.98694 | 31.6665 | 0.98601 | -0.22004 | 31.6665 | -2.2185 | -10.9182 | -6.968 |  | -1.082 | -5.886 | -0.0047688
2022-06-19 17:00 | 10.0001 | 0.016859 | 36.5812 | 0.98694 | 31.7083 | 1.0026 | -0.22014 | 31.7083 | -2.2076 | -10.7593 | -6.9801 |  | -1.0769 | -5.9033 | -0.11388
2022-06-19 18:00 | 10.0001 | 0.03154 | 36.4558 | 0.98694 | 31.7501 | 1.0026 | -0.22208 | 31.7501 | -2.333 | -12.6982 | -7.0511 |  | -1.1365 | -5.9145 | -0.079048
2022-06-19 19:01 | 10.0001 | 0.0292 | 36.3072 | 0.98694 | 31.7918 | 1.0026 | -0.22392 | 31.7918 | -2.4816 | -15.2831 | -7.1189 |  | -1.2069 | -5.912 | -0.063204
2022-06-19 20:01 | 10.0001 | 0.011921 | 36.5701 | 0.98694 | 31.8336 | 1.0026 | -0.22096 | 31.8336 | -2.2188 | -10.9227 | -7.0338 |  | -1.0822 | -5.9517 | -0.064702
2022-06-19 21:01 | 10.0001 | 0.023842 | 36.5455 | 0.98694 | 31.8754 | 1.0026 | -0.22095 | 31.8754 | -2.2433 | -11.2891 | -7.043 |  | -1.0939 | -5.9491 | -0.085204
2022-06-19 22:00 | 10.0001 | 0.026656 | 36.3335 | 0.98694 | 31.9165 | 0.98597 | -0.22194 | 31.9165 | -2.4553 | -14.8013 | -7.0835 |  | -1.1945 | -5.889 | 0.0079781
2022-06-19 23:00 | 10.0001 | 0.0292 | 35.9789 | 0.98694 | 31.9583 | 1.0026 | -0.22938 | 31.9583 | -2.81 | -22.1872 | -7.3307 |  | -1.3612 | -5.9695 | -0.059051
2022-06-20 | 10.0001 | 0.016859 | 35.6641 | 0.98694 | 32 | 1.0026 | -0.23273 | 32 | -3.1248 | -30.5109 | -7.4472 |  | -1.5071 | -5.9401 | -0.10498
2022-06-20 01:01 | 10.0001 | 0.023842 | 35.292 | 0.98694 | 32.0418 | 1.0026 | -0.23868 | 32.0418 | -3.4968 | -42.7575 | -7.6479 |  | -1.677 | -5.9709 | -0.1169
2022-06-20 02:01 | 10.0001 | 0.020648 | 35.1575 | 0.98694 | 32.0836 | 1.0026 | -0.23853 | 32.0836 | -3.6313 | -47.8839 | -7.6529 |  | -1.7376 | -5.9154 | -0.045126
2022-06-20 03:01 | 10.0001 | 0.016859 | 34.9777 | 0.98694 | 32.1254 | 1.0026 | -0.24078 | 32.1254 | -3.8111 | -55.3538 | -7.735 |  | -1.8179 | -5.9171 | -0.019977
2022-06-20 04:00 | 10.0001 | 0.023842 | 34.7145 | 0.98694 | 32.1664 | 0.98595 | -0.24595 | 32.1664 | -4.0743 | -67.6349 | -7.9113 |  | -1.934 | -5.9773 | -0.063882
2022-06-20 05:00 | 10.0001 | 0.011921 | 34.2992 | 0.98694 | 32.2082 | 1.0027 | -0.24937 | 32.2082 | -4.4897 | -90.4987 | -8.0319 |  | -2.1133 | -5.9186 | -0.0084908
2022-06-20 06:00 | 10.0001 | 0.026656 | 34.2216 | 0.98694 | 32.25 | 1.0026 | -0.25291 | 32.25 | -4.5672 | -95.2691 | -8.1565 |  | -2.1462 | -6.0103 | -0.15514
2022-06-20 07:00 | 10.0001 | NaN | 34.6656 | 0.98694 | 32.2918 | 1.0027 | -0.24609 | 32.2918 | -4.1232 | -70.0976 | -7.9468 |  | -1.9554 | -5.9914 | -0.13051
2022-06-20 08:01 | 10.0001 | 0.026656 | 34.9527 | 0.98694 | 32.3335 | 1.0026 | -0.2407 | 32.3335 | -3.8361 | -56.4501 | -7.7826 |  | -1.829 | -5.9536 | -0.085072
2022-06-20 09:01 | 10.0001 | 0.023842 | 35.0575 | 0.98694 | 32.3753 | 1.0026 | -0.23917 | 32.3753 | -3.7314 | -51.9521 | -7.7431 |  | -1.7824 | -5.9607 | -0.10826
2022-06-20 10:01 | 10.0001 | 0.016859 | 35.15 | 0.98694 | 32.4171 | 1.0026 | -0.23736 | 32.4171 | -3.6388 | -48.1816 | -7.6945 |  | -1.7409 | -5.9535 | -0.062758
2022-06-20 11:00 | 10.0001 | 0.026656 | 35.2756 | 0.98694 | 32.4582 | 0.98593 | -0.2365 | 32.4582 | -3.5132 | -43.3634 | -7.6762 |  | -1.6844 | -5.9918 | -0.054624
2022-06-20 12:00 | 10.0001 | NaN | 35.3862 | 0.98694 | 32.4999 | 1.0026 | -0.2341 | 32.4999 | -3.4027 | -39.3963 | -7.6083 |  | -1.6343 | -5.9741 | -0.083163
2022-06-20 13:00 | 10.0001 | 0.011921 | 35.5776 | 0.98694 | 32.5417 | 1.0026 | -0.23106 | 32.5417 | -3.2112 | -33.1126 | -7.5192 |  | -1.5469 | -5.9723 | -0.10146
2022-06-20 14:01 | 10.0001 | 0.016859 | 35.7538 | 0.98694 | 32.5835 | 1.0027 | -0.2294 | 32.5835 | -3.035 | -27.9562 | -7.4746 |  | -1.4657 | -6.0089 | -0.088434
2022-06-20 15:01 | 10.0001 | 0.03154 | 36.024 | 0.98694 | 32.6252 | 1.0001 | -0.22674 | 32.6252 | -2.7648 | -21.1355 | -7.3976 |  | -1.3401 | -6.0575 | -0.14536
2022-06-20 16:01 | 10.0001 | 0.023842 | 36.3095 | 0.98694 | 32.667 | 1.0026 | -0.22233 | 32.667 | -2.4794 | -15.2414 | -7.2628 |  | -1.2059 | -6.0569 | -0.17091
2022-06-20 17:01 | 10.0001 | 0.016859 | 36.6809 | 0.98694 | 32.7087 | 1.0026 | -0.21604 | 32.7087 | -2.1079 | -9.3657 | -7.0663 |  | -1.0292 | -6.037 | -0.13371
2022-06-20 18:00 | 10.0001 | 0.033717 | 36.9395 | 0.98694 | 32.7498 | 0.98593 | -0.21276 | 32.7498 | -1.8493 | -6.3243 | -6.968 |  | -0.90507 | -6.0629 | -0.14841
2022-06-20 19:00 | 10.0001 | 0.020648 | 37.1083 | 0.98694 | 32.7916 | 1.0027 | -0.20903 | 32.7916 | -1.6805 | -4.7463 | -6.8545 |  | -0.8236 | -6.0309 | -0.11892
2022-06-20 20:00 | 10.0001 | 0.023842 | 37.1888 | 0.98694 | 32.8334 | 1.0026 | -0.20722 | 32.8334 | -1.6 | -4.096 | -6.8038 |  | -0.7846 | -6.0192 | -0.067575
2022-06-20 21:01 | 10.0001 | 0.033717 | 37.1411 | 0.98694 | 32.8751 | 1.0027 | -0.20973 | 32.8751 | -1.6477 | -4.4737 | -6.895 |  | -0.80773 | -6.0873 | -0.13823
2022-06-20 22:01 | 10.0001 | 0.020648 | 36.8919 | 0.98694 | 32.9169 | 1.0026 | -0.21332 | 32.9169 | -1.897 | -6.8261 | -7.0217 |  | -0.92802 | -6.0937 | -0.20457
2022-06-20 23:01 | 10.0001 | 0.016859 | 36.7471 | 0.98694 | 32.9587 | 1.0026 | -0.21572 | 32.9587 | -2.0418 | -8.5117 | -7.1098 |  | -0.99756 | -6.1123 | -0.14267
2022-06-21 | 10.0001 | 0.023842 | 36.7717 | 0.98694 | 32.9998 | 0.98596 | -0.21533 | 32.9998 | -2.0171 | -8.2074 | -7.1058 |  | -0.98575 | -6.12 | -0.17998
2022-06-21 01:00 | 10.0001 | 0.035763 | 36.8518 | 0.98694 | 33.0415 | 1.0026 | -0.21309 | 33.0415 | -1.937 | -7.268 | -7.0409 |  | -0.9473 | -6.0936 | -0.12276
2022-06-21 02:00 | 10.0001 | 0.023842 | 37.1087 | 0.98694 | 33.0833 | 1.0026 | -0.21163 | 33.0833 | -1.6802 | -4.7431 | -7.0014 |  | -0.82342 | -6.178 | -0.26271
2022-06-21 03:00 | 10.0001 | 0.011921 | 37.3206 | 0.98694 | 33.1251 | 1.0027 | -0.20745 | 33.1251 | -1.4682 | -3.1648 | -6.8717 |  | -0.72063 | -6.1511 | -0.23405
2022-06-21 04:01 | 10.0001 | 0.03154 | 37.4118 | 0.98694 | 33.1669 | 1.0026 | -0.20548 | 33.1669 | -1.3771 | -2.6113 | -6.815 |  | -0.6763 | -6.1387 | -0.16132
2022-06-21 05:01 | 10.0001 | 0.026656 | 37.477 | 0.98694 | 33.2087 | 1.0027 | -0.20412 | 33.2087 | -1.3118 | -2.2575 | -6.7785 |  | -0.64452 | -6.134 | -0.21532
2022-06-21 06:01 | 10.0001 | 0.016859 | 37.5051 | 0.98694 | 33.2504 | 1.0026 | -0.20654 | 33.2504 | -1.2838 | -2.1157 | -6.8677 |  | -0.63084 | -6.2368 | -0.22645
2022-06-21 07:00 | 10.0001 | NaN | 37.5267 | 0.98694 | 33.2915 | 0.9859 | -0.20449 | 33.2915 | -1.2621 | -2.0106 | -6.8079 |  | -0.62029 | -6.1876 | -0.19624
2022-06-21 08:00 | 10.0001 | 0.020648 | 37.3206 | 0.98694 | 33.3333 | 1.0026 | -0.2074 | 33.3333 | -1.4682 | -3.1648 | -6.9133 |  | -0.72063 | -6.1926 | -0.23907
2022-06-21 09:00 | 10.0001 | 0.0292 | 37.28 | 0.98694 | 33.3751 | 1.0026 | -0.20778 | 33.3751 | -1.5088 | -3.4348 | -6.9345 |  | -0.74036 | -6.1942 | -0.23351
2022-06-21 10:01 | 10.0001 | 0.023842 | 37.3466 | 0.98694 | 33.4168 | 1.0026 | -0.20703 | 33.4168 | -1.4422 | -2.9998 | -6.9183 |  | -0.708 | -6.2103 | -0.25687
2022-06-21 11:01 | 10.0001 | 0.03154 | 37.6565 | 0.98694 | 33.4586 | 1.0026 | -0.20278 | 33.4586 | -1.1324 | -1.452 | -6.7846 |  | -0.5569 | -6.2277 | -0.23574
2022-06-21 12:01 | 10.0001 | NaN | 37.7019 | 0.98694 | 33.5004 | 1.0026 | -0.20107 | 33.5004 | -1.0869 | -1.2842 | -6.7359 |  | -0.53469 | -6.2013 | -0.2271
2022-06-21 13:00 | 10.0001 | NaN | 37.6686 | 0.98694 | 33.5415 | 0.98592 | -0.20104 | 33.5415 | -1.1202 | -1.4057 | -6.743 |  | -0.55095 | -6.1921 | -0.21983
2022-06-21 14:00 | 10.0001 | 0.016859 | 37.8074 | 0.98694 | 33.5832 | 1.0026 | -0.19831 | 33.5832 | -0.98147 | -0.94544 | -6.66 |  | -0.48304 | -6.1769 | -0.16809
2022-06-21 15:00 | 10.0001 | 0.03154 | 37.8863 | 0.98694 | 33.625 | 1.0027 | -0.19785 | 33.625 | -0.90252 | -0.73514 | -6.6529 |  | -0.44434 | -6.2085 | -0.15106
2022-06-21 16:01 | 10.0001 | 0.0292 | 37.8803 | 0.98694 | 33.6668 | 1.0026 | -0.19734 | 33.6668 | -0.90852 | -0.74989 | -6.6437 |  | -0.44728 | -6.1965 | -0.13954
2022-06-21 17:01 | 10.0001 | 0.020648 | 37.9863 | 0.98694 | 33.7086 | 1.0027 | -0.19637 | 33.7086 | -0.80253 | -0.51686 | -6.6194 |  | -0.39526 | -6.2242 | -0.18715
2022-06-21 18:01 | 10.0001 | 0.03154 | 38.1047 | 0.98694 | 33.7504 | 1.0026 | -0.19436 | 33.7504 | -0.68408 | -0.32012 | -6.5596 |  | -0.33706 | -6.2226 | -0.15959
2022-06-21 19:00 | 10.0001 | 0.023842 | 38.0303 | 0.98694 | 33.7914 | 0.98593 | -0.1955 | 33.7914 | -0.75855 | -0.43646 | -6.6063 |  | -0.37366 | -6.2326 | -0.20174
2022-06-21 20:00 | 10.0001 | 0.033717 | 38.0217 | 0.98694 | 33.8332 | 1.0026 | -0.19601 | 33.8332 | -0.76709 | -0.45138 | -6.6316 |  | -0.37786 | -6.2537 | -0.23453
2022-06-21 21:00 | 10.0001 | 0.026656 | 38.0286 | 0.98694 | 33.875 | 1.0026 | -0.19574 | 33.875 | -0.76019 | -0.4393 | -6.6306 |  | -0.37446 | -6.2561 | -0.16884
2022-06-21 22:00 | 10.0001 | 0.020648 | 38.1342 | 0.98694 | 33.9168 | 1.0026 | -0.19439 | 33.9168 | -0.65458 | -0.28047 | -6.5931 |  | -0.32255 | -6.2705 | -0.1769
2022-06-21 23:01 | 10.0001 | 0.011921 | 37.9013 | 0.98694 | 33.9585 | 1.0027 | -0.19821 | 33.9585 | -0.8875 | -0.69906 | -6.7309 |  | -0.43697 | -6.2939 | -0.18167
2022-06-22 00:01 | 10.0001 | 0.020648 | 37.9277 | 0.98694 | 34.0003 | 1.0026 | -0.19785 | 34.0003 | -0.86109 | -0.63847 | -6.7268 |  | -0.42401 | -6.3028 | -0.1827
2022-06-22 01:01 | 10.0001 | 0.026656 | 37.9224 | 0.98694 | 34.0421 | 1.0026 | -0.19757 | 34.0421 | -0.86641 | -0.65038 | -6.7258 |  | -0.42662 | -6.2992 | -0.20545
2022-06-22 02:00 | 10.0001 | 0.041295 | 37.8058 | 0.98694 | 34.0832 | 0.98593 | -0.19888 | 34.0832 | -0.98305 | -0.95002 | -6.7785 |  | -0.48382 | -6.2947 | -0.11671
2022-06-22 03:00 | 10.0001 | 0.023842 | 37.5672 | 0.98694 | 34.1249 | 1.0026 | -0.20303 | 34.1249 | -1.2216 | -1.8231 | -6.9285 |  | -0.60051 | -6.328 | -0.17689
2022-06-22 04:00 | 10.0001 | 0 | 37.7139 | 0.98694 | 34.1667 | 1.0026 | -0.20068 | 34.1667 | -1.0749 | -1.2421 | -6.8565 |  | -0.52881 | -6.3277 | -0.18906
2022-06-22 05:01 | 10.0001 | 0.035763 | 37.4714 | 0.98694 | 34.2085 | 1.0026 | -0.20378 | 34.2085 | -1.3174 | -2.2864 | -6.971 |  | -0.64723 | -6.3238 | -0.18983
2022-06-22 06:01 | 10.0001 | 0.011921 | 37.2174 | 0.98694 | 34.2503 | 1.0026 | -0.20664 | 34.2503 | -1.5714 | -3.8804 | -7.0774 |  | -0.77074 | -6.3067 | -0.069842
2022-06-22 07:01 | 10.0001 | 0.0292 | 37.2139 | 0.98694 | 34.292 | 1.0026 | -0.2076 | 34.292 | -1.5749 | -3.9065 | -7.1189 |  | -0.77245 | -6.3465 | -0.15881
2022-06-22 08:00 | 10.0001 | 0.033717 | 37.2142 | 0.98694 | 34.3331 | 0.98593 | -0.20702 | 34.3331 | -1.5746 | -3.9043 | -7.1078 |  | -0.77231 | -6.3355 | -0.14289
2022-06-22 09:00 | 10.0001 | 0.020648 | 37.4205 | 0.98694 | 34.3749 | 1.0026 | -0.20368 | 34.3749 | -1.3684 | -2.5622 | -7.0014 |  | -0.67207 | -6.3293 | -0.13519
2022-06-22 10:00 | 10.0001 | NaN | 37.6662 | 0.98694 | 34.4167 | 1.0027 | -0.19922 | 34.4167 | -1.1226 | -1.4148 | -6.8565 |  | -0.55213 | -6.3044 | -0.09408
2022-06-22 11:01 | 10.0001 | 0.011921 | 37.9707 | 0.98694 | 34.4585 | 1.0027 | -0.19657 | 34.4585 | -0.81815 | -0.54764 | -6.7734 |  | -0.40293 | -6.3705 | -0.14285
2022-06-22 12:01 | 10.0001 | 0.020648 | 38.0803 | 0.98694 | 34.5002 | 1.0026 | -0.19407 | 34.5002 | -0.70848 | -0.35562 | -6.6954 |  | -0.34905 | -6.3464 | -0.14513
2022-06-22 13:01 | 10.0001 | 0.033717 | 38.2327 | 0.98694 | 34.542 | 1.0027 | -0.19295 | 34.542 | -0.55609 | -0.17197 | -6.665 |  | -0.27409 | -6.3909 | -0.19873
2022-06-22 14:00 | 10.0001 | 0.026656 | 38.5009 | 0.98694 | 34.5831 | 0.98598 | -0.18827 | 34.5831 | -0.28797 | -0.023881 | -6.511 |  | -0.14201 | -6.369 | -0.19212
2022-06-22 15:00 | 10.0001 | 0.0292 | 38.6582 | 0.98694 | 34.6249 | 1.0026 | -0.18611 | 34.6249 | -0.13063 | -0.002229 | -6.4441 |  | -0.064428 | -6.3797 | -0.17121
2022-06-22 16:00 | 10.0001 | 0.020648 | 38.5543 | 0.98694 | 34.6666 | 1.0026 | -0.18571 | 34.6666 | -0.23451 | -0.012898 | -6.4381 |  | -0.11566 | -6.3224 | -0.12595
2022-06-22 17:00 | 10.0001 | 0.016859 | 38.53 | 0.98694 | 34.7084 | 1.0026 | -0.18689 | 34.7084 | -0.25884 | -0.017342 | -6.4867 |  | -0.12765 | -6.359 | -0.1349
2022-06-22 18:01 | 10.0001 | 0.023842 | 38.5171 | 0.98694 | 34.7502 | 1.0026 | -0.18757 | 34.7502 | -0.27176 | -0.02007 | -6.5181 |  | -0.13402 | -6.3841 | -0.16152
2022-06-22 19:01 | 10.0001 | 0.033717 | 38.494 | 0.98694 | 34.792 | 1.0026 | -0.1865 | 34.792 | -0.2948 | -0.025621 | -6.4887 |  | -0.14538 | -6.3433 | -0.11068
2022-06-22 20:01 | 10.0001 | 0.020648 | 38.5686 | 0.98694 | 34.8337 | 1.0026 | -0.18744 | 34.8337 | -0.22026 | -0.010686 | -6.5292 |  | -0.10863 | -6.4206 | -0.1668
2022-06-22 21:00 | 10.0001 | 0.020648 | 38.5086 | 0.98694 | 34.8748 | 0.98601 | -0.18797 | 34.8748 | -0.28024 | -0.022008 | -6.5556 |  | -0.1382 | -6.4174 | -0.16131
2022-06-22 22:00 | 10.0001 | 0 | 38.2797 | 0.98694 | 34.9166 | 1.0026 | -0.19051 | 34.9166 | -0.50911 | -0.13196 | -6.6518 |  | -0.25096 | -6.4009 | -0.13037
2022-06-22 23:00 | 10.0001 | 0.03154 | 38.5647 | 0.98694 | 34.9584 | 1.0026 | -0.18712 | 34.9584 | -0.22417 | -0.011264 | -6.5414 |  | -0.11055 | -6.4309 | -0.13696
2022-06-23 00:01 | 10.0001 | 0.026656 | 38.8438 | 0.98694 | 35.0002 | 1.0026 | -0.18276 | 35.0002 | 0.054992 | 0.0001663 | -6.3965 |  | 0.027124 | -6.4236 | -0.12084
2022-06-23 01:01 | 10.0001 | 0.0292 | 38.8555 | 0.98694 | 35.0419 | 1.0026 | -0.18268 | 35.0419 | 0.066631 | 0.00029582 | -6.4016 |  | 0.032865 | -6.4344 | -0.13527
2022-06-23 02:01 | 10.0001 | 0.023842 | 38.9006 | 0.98694 | 35.0837 | 1.0027 | -0.18252 | 35.0837 | 0.11182 | 0.0013982 | -6.4036 |  | 0.055153 | -6.4588 | -0.16399
2022-06-23 03:00 | 10.0001 | 0 | 38.8308 | 0.98694 | 35.1248 | 0.98595 | -0.1839 | 35.1248 | 0.041934 | 7.3737e-05 | -6.4593 |  | 0.020683 | -6.48 | -0.15209
2022-06-23 04:00 | 10.0001 | 0.026656 | 38.6869 | 0.98694 | 35.1666 | 1.0026 | -0.18489 | 35.1666 | -0.10197 | -0.0010602 | -6.5019 |  | -0.050293 | -6.4516 | -0.12391
2022-06-23 05:00 | 10.0001 | 0.03154 | 38.6529 | 0.98694 | 35.2083 | 1.0026 | -0.1853 | 35.2083 | -0.13597 | -0.0025136 | -6.5242 |  | -0.067061 | -6.4571 | -0.13336
2022-06-23 06:01 | 10.0001 | 0 | 38.6023 | 0.98694 | 35.2501 | 1.0026 | -0.18675 | 35.2501 | -0.18648 | -0.006485 | -6.5829 |  | -0.091973 | -6.491 | -0.18434
2022-06-23 07:01 | 10.0001 | 0.016859 | 38.5683 | 0.98694 | 35.2919 | 1.0026 | -0.18607 | 35.2919 | -0.22057 | -0.010731 | -6.5667 |  | -0.10878 | -6.458 | -0.11148
2022-06-23 08:01 | 10.0001 | 0.035763 | 38.407 | 0.98694 | 35.3337 | 1.0026 | -0.18766 | 35.3337 | -0.3818 | -0.055655 | -6.6306 |  | -0.18826 | -6.4423 | -0.10676
2022-06-23 09:01 | 10.0001 | 0.016859 | 38.6391 | 0.98694 | 35.3754 | 1.0026 | -0.1854 | 35.3754 | -0.14971 | -0.0033554 | -6.5586 |  | -0.073839 | -6.4848 | -0.15536
2022-06-23 10:00 | 10.0001 | 0.020648 | 38.6537 | 0.98694 | 35.4165 | 0.98591 | -0.18547 | 35.4165 | -0.13514 | -0.0024679 | -6.5688 |  | -0.066652 | -6.5021 | -0.19775
XXX